C.  Hydrographic Measurement Techniques and Calibration

- CTD/O2 Measurements

13 December 2013

- Personnel

Tetsuya NAKAMURA (GEMD/JMA)

Yoshikazu HIGASHI (GEMD/JMA)

Tomoyuki KITAMURA (GEMD/JMA)

Keizo SHUTTA (GEMD/JMA)

Takahiro SEGAWA (GEMD/JMA)

Yasuaki BUNGI (GEMD/JMA)

(2)  CTD Traction Winch and Motion Compensated Crane Arrangements

The CTD/O2 system was deployed by using a Traction Winch System with ca. 7000m of 8.03 mm armored cable (Tyco Electronics, USA) and a Motion Compensated Crane (Dynacon, Inc., USA). The system was installed on the R/V Ryofu Maru in March, 2010 (Photo C1.1).

Photo C1.1. (Left) The Traction Winch and (right) Motion Compensated Crane.

(3)  Overview of the CTD/O2 system

The CTD/O2 system, SBE 911plus system (Sea-Bird Electronics, Inc., USA), was used for entire cruise. The system is consisted of a SBE 9plus underwater unit and a SBE 11plus deck unit. The SBE 11plus deck unit is a rack-mountable interface which supplies DC power to under water unit, decodes serial data stream, formats data under microprocessor control, and passes the data to a computer. The real time serial data from the underwater unit is sent to the deck unit. The deck unit decodes the serial data and sends them to a personal computer to display and a storage in a file using SEASAVE data acquisition software (SEASAVE-Win32, version 7.18).

The SBE 911plus system controls 36-position SBE 32 Carousel Water Sampler (Photo C1.2). The Carousel with a custom frame accepts 10-liter Niskin bottles (General Oceanics, Inc., USA). The SBE 9plus was mounted horizontally in the 36-position carousel frame. Two set of SBE's temperature (SBE 3plus) and conductivity (SBE 4C) sensor modules were used with the SBE 9plus underwater unit. Two modular units of underwater housing pump (SBE 5T) flush water through sensor tubing at a constant rate independent of the CTD's motion (Photo C1.3). Two dissolved oxygen sensors (RINKO III: JFE Advantech Co., Ltd., Japan; http://www.jfe-alec.co.jp/html/english_top.htm) were mounted on CTD housing, by the side of primary T/C sensors (Photo C.1.3). Auxiliary sensors, Deep Ocean Standards Thermometer (SBE 35) and an altimeter (PSA-916D: Teledyne Benthos, Inc., USA) were also used with the SBE 9plus underwater unit. The SBE 35 was mounted at the center of CTD between two pumps and the altimeter was mounted at the same height of pressure sensor of SBE 9plus.

Photo C1.2. The CTD/O2 system (left) top view and (right) bottom view.

SBE35

Photo C1.3.  (left) SBE 9plus CTD with SBE35 and (right) RINKO III.

Table C.1.1. Specification and serial number of the CTD/O2 measurements system components.

| Deck unit | Serial Number |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SBE 11plus （SBE) | 0648 |  |  |  |  |  |  |  |
| Under water unit | Serial Number |  | Range |  | Accuracy |  | Stability | Resolution |
| SBE 9plus (SBE) | 35560 (Pressure: 0764) |  | 0 to 10000 psi 0 to 6800 dbar |  | 0.015 %(FS) 1.0 dbar |  | 0.002%FS/year 0.2 dbar/year | 0.001 % (FS) 0.1 dbar |
| Temperature | Serial Number |  | Range |  | Accuracy |  | Stability | Resolution |
| SBE 3plus (SBE) | 4923 (primary) 4199 (secondary) |  | -5 to 35 ℃ |  | 0.001 ℃ |  | 0.0002 ℃/month | 0.0002 ℃ |
| Conductivity | Serial Number |  | Range |  | Accuracy |  | Stability | Resolution |
| SBE 4C (SBE) | 3670 (primary) 2842 (secondary) |  | 0 to 7 S/m |  | 0.0003 S/m |  | 0.0003 S/m/month | 0.00004 S/m |
| Pump | Serial Number |  |  |  |  |  |  |  |
| SBE 5T (SBE) | 3887 (primary) 5501 (secondary) |  |  |  |  |  |  |  |
| Oxygen | Serial Number |  | Range |  | Linearity |  | Response time | Resolution |
| RINKO III (JFE) | 25 (primary, foil number:144) 26 (secondary, foil numner:144) |  | 0 to 200% (saturation) |  | ±2% (FS) |  | ≤ 1 second | - to 0.04 % |
| Water sampler | Serial Number |  |  |  |  |  |  |  |
| SBE 32 (SBE) | 0734 |  |  |  |  |  |  |  |
| Altimeter | Serial Number |  | Range |  | Resolution |  |  |  |
| PSA-916D (TB) | 1267 |  | 0 to 100 m |  | 1 cm |  |  |  |
| Water Sampling Bottle |  |  |  |  |  |  |  |  |
| Niskin Bottle (GO) | ･10-Liter ･No TEFRON coating |  |  | ･Bottle O-ring: Viton ･Stainless spring |  |  |  |  |

SBE: Sea-Bird Electronics Inc., USA	JFE: JFE Advantech Co., Ltd., Japan

GO: General Oceanics, Inc., USA	TB: Teledyne Benthos, Inc., USA

(4)  Pre-cruise calibration

(4.1) Pressure

Pre-cruise calibration were performed at SBE, Inc., USA. The following coefficients were used in the SEASOFT:

S/N 0764, 25 May 2010

| c1 | = | –4.318853e+04 |
| --- | --- | --- |
| c2 | = | –4.853949e–01 |
| c3 | = | 1.294200e–02 |
| d1 | = | 3.706500e–02 |
| d2 | = | 0.000000e+00 |
| t1 | = | 3.005385e+01 |
| t2 | = | –4.407111e–04 |
| t3 | = | 4.098190e–06 |
| t4 | = | 1.662250e–09 |
| t5 | = | 0.000000e+00 |

Pressure coefficients are first formulated into

where U is temperature in degrees Celsius. The pressure temperature, U, is determined according to

The following coefficients were used for S/N 0764 in SEASOFT:

M=1.289080e–02

B=–8.282450e+00

(in the underwater unit system configuration sheet dated on 25 May, 2010)

Finally, pressure is computed as

where t is pressure period (μsec).

Since the pressure sensor measures the absolute value, it inherently includes atmospheric pressure (about 14.7 psi). SEASOFT subtracts 14.7 psi from computed pressure above automatically.

The pressure sensor drift is known to be primarily an offset drift at all pressures rather than a change of span slope. The following coefficients for the sensor drift correction were also used in SEASOFT:

Slope=0.999930

Offset=–0.56680

The drift–corrected pressure is computed as

(4.2) Temperature (SBE 3plus)

Pre-cruise calibrations were performed at SBE, Inc., USA. The following coefficients were used in SEASOFT:

S/N 4923(primary), 26 May 2010

| g | = | 4.35306322e–03 |
| --- | --- | --- |
| h | = | 6.39215989e–04 |
| i | = | 2.11728148e–05 |
| j | = | 1.77647263e–06 |
| f0 | = | 1000.000 |

S/N 4199(secondary), 26 May 2010

| g | = | 4.39450115e–03 |
| --- | --- | --- |
| h | = | 6.49623486e–04 |
| i | = | 2.38724882e–05 |
| j | = | 2.21735485e–06 |
| f0 | = | 1000.000 |

Temperature (ITS-90) is computed according to

where f is the instrument frequency (Hz).

Time drift of the SBE 3plus temperature sensors based on the laboratory calibrations is shown in Figure C.1.1.

Figure C.1.1. Time drift of the SBE 3plus temperature sensors (S/N 4923 and 4199) based on laboratory calibrations performed by SBE, Inc. The secondary sensor (S/N4199) was resecured the temperature probe retaining nut in December 2007, and replaced the main piston O-rings in August 2009

(4.3) Conductivity (SBE 4C)

Pre-cruise sensor calibrations were performed at SBE, Inc., USA. The following coefficients were used in SEASOFT:

S/N 3670(primary), 26 May 2010

| g | = | –1.02022781e+001 |
| --- | --- | --- |
| h | = | 1.57745207e+000 |
| i | = | –2.48735605e–003 |
| j | = | 2.86313468e–004 |
| CPcor | = | –9.57e–08 |
| CTcor | = | 3.25e–06 |

S/N 2842(secondary), 26 May 2010

| g | = | –1.01321263e+001 |
| --- | --- | --- |
| h | = | 1.38952824e+000 |
| i | = | 2.52094473e–004 |
| j | = | 4.58018677e–005 |
| CPcor | = | –9.57e–08 |
| CTcor | = | 3.25e–06 |

Conductivity of a fluid in the cell is expressed as:

where f is the instrument frequency (kHz), t is the water temperature (degrees Celsius) and p is the water pressure (dbar).

(4.4) Deep Ocean Standards Thermometer (SBE 35)

In the first place a newly manufactured SBE 35 is first calibrated in a temperature controlled bath against Standard Platinum Resistance Thermometer, and this calibration is referred as the Linearization Calibration. In the next place SBE 35 is calibrated to generate slope and offset coefficients that correct for the time drift from the Linearization Calibration. This calibration is referred Fixed Point Calibrations. Pre-cruise sensor calibrations were performed at SBE, Inc., USA. The following coefficients were stored in EEPROM:

S/N 0069, 23 October, 2006(1st step: Linearization Calibration)

| a0 | = | 4.96812728e–003 |
| --- | --- | --- |
| a1 | = | –1.39341438e–003 |
| a2 | = | 2.06596098e–004 |
| a3 | = | –1.14827915e–005 |
| a4 | = | 2.44200422e–007 |

Linearized temperature (ITS–90) is computed according to

where n is the instrument output. Then the SBE 35 is certified by measurements in thermodynamic fixed-point cells of the Triple Point of Water (TPW: 0.0100 degrees Celsius) and Gallium Melt Point (GaMP: 29.7646 degrees Celsius). The slow time drift of the SBE 35 is adjusted by periodic recertification corrections.

S/N 0069, 21 August, 2009 (2nd step: Fixed Point Calibration)

Slope=0.999998

Offset=0.000258

Temperature (ITS-90) is calibrated according to

The time required per sample = 1.1 * NCYCLES + 2.7 seconds. The 1.1 seconds is total time per an acquisition cycle. NCYCLES is the number of acquisition cycles per sample. The 2.7 seconds is required for converting the measured values to temperature and storing average in EEPROM. In this cruise NCYCLES was set to 2.

(5) Data processing

(5.1) Data Collection

CTD system was powered on at least five minutes in advance of the operation and was powered off after CTD came up from the surface.

The package was lowered into the water from the port side and held about 10 m beneath the surface for about one minute in order to activate the pump. After the pump was activated, the package was lifted to the surface and lowered at a rate of 0.6 m/s approximately to 50m (or more when wave height was high), then the package was stopped to turn on the heave compensator of the crane. The package was lowered again at a rate of 0.9 m/s to the bottom. For the up cast, the package was lifted at a rate of 0.9 m/s except for bottle firing stops. At each bottle firing stops, the bottle was fired after waiting for about 30 seconds and the package was stayed at least 10 seconds for measurement of the SBE 35 after firing. At 50 m from the surface, the package was stopped to turn off the heave compensator of the crane.

Water samples were collected using a 36-position SBE 32 Carousel Water Sampler with 10-liter Niskin bottles.

The SBE11plus deck unit received the data signal from the CTD. Digitized data were forwarded to a personal computer running the SEASAVE data acquisition software (SEASAVE-Win32, version 7.18). Temperature, conductivity, salinity, oxygen and descent/ascent rate profiles were displayed in real–time with the package depth, altimeter reading and sound speed. Differences in temperature, salinity, and oxygen between primary and secondary sensor were also displayed in order to monitor the status of sensors. Note that oxygen data were displayed and monitored in voltage (0–5V).

Altimeter (PSA-916D) was mounted at the same height of pressure sensor of SBE 9plus (Photo C1.4). The altimeter detected the sea floor at 100 of 124 stations, the average distance of beginning detecting the sea floor was 26.4m, and that of final detection of sea floor was 13.8m. The summary of detection of PSA-916D was shown in Figure C.1.2.

Photo C.1.4. The location of PSA-916D.

Figure C.1.2. The summary of detection of PSA-916D. The left panel shows the stations of detection, the right panel shows the relationship among PSA-916D, bathymetry and CTD depth.

(5.2) Data Processing

SEASOFT (SEASOFT-Win32, version 7.18) consists of modular menu driven routines for acquisition, display, processing, and archiving of oceanographic data acquired with SBE equipment, and is designed to work with a personal computer. Raw data are acquired from instruments and are stored as unmodified data. The conversion module DATCNV uses instrument configuration and calibration coefficients to create a converted engineering unit data file that is operated on by all SEASOFT post processing modules.

Each SEASOFT module that modifies the converted data file adds proper information to the header of the converted file permitting tracking of how the various oceanographic parameters were obtained. The converted data is stored in rows and columns of ASCII numbers. The last data column is a flag field used to mark scans as good or bad.

We made the original module for the process of RINKO III and JMA’s report. The following are the SEASOFT data processing module and JMA original module sequence and specifications used in the reduction of CTD data in this cruise.

DATCNV converted the raw data to engineering unit data such as scan number, pressure, temperatures, conductivities, RINKO III voltages, time in Julian days, pump status, and flag. DATCNV also extracted bottle information where scans were marked with the bottle confirm bit during acquisition. The duration was set to 2.0 seconds, and the offset was set to 0.0 seconds.

DECKP_OFF (original module) cancelled the deck pressure and after this module, spikes in temperature and salinity were eliminated manually.

RINKO_hystoff (original module) cancelled the hysteresis of RINKO III using the method of SBE 43 (Sea-Bird Electronics, 2009) .

SECT_IN (original module) found the first and last scan numbers while pump was activated, and made the surface data while pump was not activated for down cast.

SECTION selected a time span of data based on scan number in order to reduce a file size. The minimum number was set to be the start time when the CTD package was beneath the sea-surface after activation of the pump. The maximum number was set to be the end time when the package came up from the surface.

FILTER performed a low pass filter on pressure with a time constant of 0.15 seconds. In order to produce zero phase lag (no time shift) the filter runs forward first then backwards.

ALIGNCTD converted the time-sequence of RINKO III sensor outputs into the pressure sequence to ensure that all calculations were made using measurements from the same parcel of water. RINKO III sensor output delays 1 second compared to pressure, temperature and conductivity.

ALIGNROS (original module) replace the RINKO III output of the bottle to that of all scan data applied ALIGNCTD module.

BOTTOLESUM created a summary of the bottle data. The bottle position, date, time were output as the first two columns. Salinities, pressure, temperatures, conductivities and oxygen voltage were averaged over 2.0 seconds.

CELLTM used a recursive filter to remove conductivity cell thermal mass effects from the measured conductivity. Typical values used were thermal anomaly amplitude alpha = 0.03 and the time constant 1/beta = 7.0.

LOOPEDIT marked scans where the CTD was moving less than the minimum velocity of 0.0 m/s (traveling backwards due to ship roll).

BINAVG averaged the data into 1dbar pressure bins. The center value of the first bin was set equal to the bin size. The bin minimum and maximum values are the center value plus and minus half the bin size. Scans with pressures greater than the minimum and less than or equal to the maximum were averaged. Scans were interpolated so that a data record could exist in every dbar.

RSC2ASC (original module) made the data set from 1dbar to the bottom of observation.

The RINKO III processes (original module) to make down and up cast data in every dbar were performed after those processes.

(6) Post–cruise calibration

(6.1) Pressure

The CTD pressure sensor offset in the period of this cruise is estimated from the pressure readings on the ship deck. In order to get the calibration data for the pre-cast pressure sensor drift, the CTD deck pressure was averaged over five scan pressure data after the CTD system had been stable on the deck.

Deck pressure was used to cancel the CTD pressure sensor offset in CTD data processing. Time series of the CTD deck pressure is shown in Figure C.1.3. Tendencies of CTD deck pressure and air pressure were almost similar during the cruise.

Figure C.1.3. Time series of the CTD deck pressure. Red line indicates atmospheric pressure anomaly. Blue line and dots indicate pre-cast deck pressure and average.

Post-cruise sensor calibrations were performed at SBE, Inc., USA. The pressure sensor drift is known to be primarily an offset drift at all pressures rather than a change of span slope.

S/N 0764, 27 September 2010

Slope = 0.999940

Offset = –0.55550

The pressure sensor drift was estimated to be 0.07 dbar at the pressure of 6000 dbar. The pressure sensor drift was small, so post-cruise calibration is not applied.

(6.2) Temperature

Budeus and Schneider (1998) noted that the CTD temperature sensor (SBE 3plus) showed a pressure sensitivity. The pressure sensitivity for a SBE 3plus sensor is usually less than +2 mK/6000 dbar. It is somewhat difficult to measure this effect in the laboratory and the difficulty is one of the primary reasons to use the SBE 35 at sea for critical work. Also SBE 3plus measurements may be affected by viscous heating (about +0.5 mK) that occurs in a TC duct and does not occur for un–pumped SBE 35 measurements (Larson and Pederson, 1996). Furthermore, the SBE 35 calibrations have some uncertainty (about 0.2 mK) and SBE 3plus calibrations have some uncertainty (about 1 mK). So the practical corrections for CTD temperature data can be made by using a SBE 35, correcting the SBE 3plus to agree with the SBE 35 (Uchida et al., 2007).

Post-cruise sensor calibration for the SBE 35 was performed at SBE, Inc., USA.

S/N 0069, 03 October 2010 (2nd step: fixed point calibration)

Slope=1.000009

Offset=0.000313

The discrepancy between the CTD temperature and the SBE 35 temperature is considered to be a function of pressure and time. But the time drift correction is regarded as 0 due to following reasons; 1) The time drift of the SBE 3plus estimated to be as –0.00094 K/year for S/N 4923, –0.00044 K/year for S/N 4199 and that of SBE 35 is estimated to be as +0.01 mK during cruise according to pre-cruise and post-cruise calibrations performed at SBE, 2).

Effect of the viscous heating is assumed to be constant. Since the pressure sensitivity is thought to be constant in time at least during observation period, the CTD temperature is calibrated as

where T is the CTD temperature in degrees Celsius, P is pressure in dbar and c0, c1 are calibration coefficients.

The calibration is performed for the primary and secondary temperature data. The CTD data created by the software module BOTTLESUM are used. (The deviation of CTD temperature from the SBE35 temperature at depth shallower than 1900 dbar is large for determining the coefficients with sufficient accuracy since the vertical temperature gradient is too large in the regions. So the coefficients are determined by least squares method using the data for the depth deeper than 1900 dbar). The temperature calibration summary is listed in Table C.1.1 at Pressure ≥ 1900dbar. We adopted secondary conductivity sensor (S/N 2842), so secondary temperature sensor (S/N4199) is adopted.

| S/N | Num | c0 (K) | c1 (K/dbar) | Average (K) | STD (K) | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 4923 | 620 | 5.2348920e–4 | 2.3591956e–7 | 0.0000 | 0.0002 | Leg 1 |
| 4923 | 378 | 5.8598217e–4 | 2.3668426e–7 | 0.0000 | 0.0002 | Leg 2 |
| 4199 | 620 | 2.0889685e–4 | 2.4549467e–7 | 0.0000 | 0.0002 | Leg 1 |
| 4199 | 378 | 2.3835674e–4 | 2.3980066e–7 | 0.0000 | 0.0002 | Leg 2 |

Table C.1.1. Temperature Calibration summary (Pressure ≥ 1900dbar).

Figure C.1.4. Difference between the CTD temperature (secondary) and the Deep Ocean Standards thermometer (SBE35) at Leg 1. Blue and red dots indicate before and after the calibration using SBE35 data respectively. Lower two panels show histogram of the difference after calibration.

Figure C.1.5. Difference between the CTD temperature (secondary) and the Deep Ocean Standards thermometer (SBE35) at Leg 2. Blue and red dots indicate before and after the calibration using SBE35 data respectively. Lower two panels show histogram of the difference after calibration.

(6.3) Salinity

The CTD salinity is computed from pressure, conductivity and temperature according to algorithm of the Practical Salinity Scale of 1978 (PSS78). The discrepancy between the CTD conductivity and the bottle conductivity is considered to be a function of pressure and time according to McTaggart et al. (2010).

Post–cruise sensor calibrations were performed in September 2010 at SBE, Inc., USA. According to the conductivity calibration report, the drifts since pre-cruise calibration was –0.00090 /month in PSS78 at 3.0 S/m for primary sensor (S/N3670), so the effect of SBE 4C drift during the cruise was estimated to be less than 0.002 in PSS78. However the time coefficient was set to zero in this cruise because the calibration was performed considering the sudden station-dependent shifts of the CTD conductivity and other calibration coefficients included the effect of slow drift by calibration grouping. So the CTD conductivity is calibrated as below.

where C is the CTD conductivity and ci and pj are calibration coefficients. Coefficient sets of each (I, J) combination was calculated by least square method between CTD conductivity and the bottle conductivity data except for bad bottle data. In calculated coefficient sets, the best (I, J) combination are determined by referring to AIC (Akaike, 1974). According to McTaggart et al. (2010), maximum of I and J are 2.

The discrepancy between the calibrated CTD conductivity and the bottle conductivity was within 0.0001 S/m for each sensor. The results of post–cruise calibration for the CTD salinity (S/N 2849) are summarized in Figure C.1.6. The calibration coefficients and the data (Num) used for the calibration are listed in Table C.1.2, and the calibration summary are listed in Table C.1.3 and C.1.4 for S/N 3670 and S/N 2842, respectively. Secondary sensor (S/N 2842) is adopted because of large time drift of primary sensor (S/N 3670).

Table C.1.2. Conductivity Calibration Coefficient Summary.

| S/N | Num | c0(mS/m) | c1 | c2(mS/m) | Stations |
| --- | --- | --- | --- | --- | --- |
|  |  |  | p1(mS/dbar) | p2(mS/m/dbar2) |  |
| 3670 | 1274 | 1.5107e–3 | –7.4144e–5 | 0.0000e+0 | Stn. 1 – 67 |
|  |  |  | 6.6856e–7 | –8.3866e–11 |  |
| 3670 | 308 | 2.2680e-3 | –8.0696e–5 | 0.0000e+0 | Stn. 68 – 83, Stn. 105 – 107 |
|  |  |  | –1.2437e–8 | 0.5038e–11 |  |
| 3670 | 608 | 1.0048e-3 | –7.6991e–5 | 0.0000e+0 | Stn. 84 – 104, Stn. 108 – 124 |
|  |  |  | 3.9031e–7 | –4.2466e–11 |  |
| 2849 | 2195 | 2.1693e-3 | –5.5359e–5 | 0.0000e+0 | Stn. 1 – 124 |
|  |  |  | 8.3709e–7 | –7.6495e–11 |  |

Table C.1.3. Conductivity Calibration Summary for S/N 3670.

| Stations | Pressure < 1900dbar |  |  | Pressure ≥ 1900 dbar |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Num | Average (mS/cm) | Std (mS/cm) | Num | Average (mS/cm) | Std (mS/cm) |
| Stn. 1 – 67 | 830 | 0.0000 | 0.0022 | 444 | –0.0000 | 0.0006 |
| Stn. 68 – 83, 105 – 107 | 211 | 0.0000 | 0.0024 | 94 | 0.0000 | 0.0004 |
| Stn. 84 – 104, 108 – 124 | 411 | 0.0000 | 0.0023 | 197 | -0.0000 | 0.0005 |

Table C.1.4. Conductivity Calibration Summary for S/N 2842.

| Stations | Pressure < 1900dbar |  |  | Pressure ≥ 1900 dbar |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Num | Average(mS/cm) | Std(mS/cm) | Num | Average(mS/cm) | Std(mS/cm) |
| Stn. 1 – 124 | 1455 | 0.0000 | 0.0022 | 740 | –0.0000 | 0.0004 |

Figure C.1.6. Difference between the CTD conductivity (secondary) and the bottle conductivity. Blue and red dots indicate before and after the calibration using bottle data respectively. Lower two panels show histogram of the difference before and after calibration.

(6.4) Oxygen

RINKO III (JFE Advantech Co., Ltd., Japan) is based on the ability of selected substance to act as dynamic fluorescence quenchers. RINKO III model is designed to use with a CTD system which accept an auxiliary analog sensor, and is designed to operate down to 7000 m. The CTD oxygen is calculated using RINKO III output (voltage) by the Stern–Volmer equation, according to a method by Uchida et al. (2008). The formulas are as follows:

Where P is the pressure in dbar, t is the potential temperature, v is RINKO output voltage in volt, T is elapsed time of the sensor from the beginning of first station in calculation group in day and [O2] is the dissolved oxygen saturation, dissolved oxygen is calculated from [O2], potential temperature and salinity by Garcìa and Gordon (1992) in μmol/kg. Calibration coefficients (c1–c9) are determined by minimizing difference between CTD oxygen and bottle dissolved oxygen by quasi-newton method (Shanno, 1970).

In general, the calibration was performed for each Leg. But in this cruise, both RINKO III had large time drift especially in the early part of Leg 1, we could not solve on c7 and c8. Worse yet, bottle dissolved oxygen data were flagged bad due to the problem of titration (please refer to C3 - (15)) in the latter part of Leg1 (Stn. 58 - 67). To avoid extrapolation in the direction of T in the period when bottle dissolved oxygen data were flagged bad, the calibration was performed across Leg1 and Leg2. The calibration was performed only for primary sensor (S/N 25) because the output of secondary sensor (S/N 26) was very noisy during the cruise. Calibration coefficients are listed in Table C.1.5. The result of the calibration during cruise is shown in Figure C.1.7, the data summary is listed in Table C.1.6 and Table C.1.7.

| Stations | c1 | c2 | c3 | c4 | c5 |
| --- | --- | --- | --- | --- | --- |
|  | c6 | c7 | c8 | c9 |  |
| Stn. 1 – 29 | 1.89977 | 1.93812e–2 | 1.43624e–4 | –6.661791e–4 | –1.26432e–1 |
|  | 3.07029e–1 | –7.05462e–5 | 9.73313e–4 | 4.45363e–2 |  |
| Stn. 30 – 80 | 1.90512 | 2.00330e–2 | 1.50096e–4 | –7.75255e–4 | –1.27249e–1 |
|  | 3.08390e–1 | –5.89936e–4 | 8.77397e–4 | 4.23961e–2 |  |
| Stn. 81 – 124 | 1.76044 | 1.55258e–2 | 8.25974e–6 | –1.65719e–3 | –9.90824e–2 |
|  | 3.06244e–1 | 1.44412e–4 | 4.13619e–4 | 4.51073e–2 |  |

Table C.1.5. Dissolved Oxygen Calibration Coefficients.

|  | Pressure < 950dbar |  |  | Pressure ≥ 950dbar |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Stations | Num | Average of deviation (μmol/kg) | STD of deviation (μmol/kg) | Num | Average of deviation (μmol/kg) | STD of deviation (μmol/kg) |
| Stn. 1 – 29 | 245 | –0.01 | 0.94 | 234 | –0.01 | 0.33 |
| Stn. 30 – 80 | 467 | 0.05 | 0.81 | 432 | 0.00 | 0.23 |
| Stn. 81 – 124 | 435 | –0.01 | 0.77 | 308 | 0.01 | 0.25 |

Table C.1.6. Dissolved Oxygen Calibration Summary.

Figure C.1.7. Difference between the CTD oxygen and bottle dissolved oxygen in the early part of Leg 1. Red dots in upper two panels indicate the result of calibration. Lower two panels show histogram of the difference between calibrated oxygen and bottle oxygen.

Figure C.1.8. Difference between the CTD oxygen and bottle dissolved oxygen across Leg 1 and Leg 2. Red dots in upper two panels indicate the result of calibration. Lower two panels show histogram of the difference between calibrated oxygen and bottle oxygen.

Figure C.1.9. Difference between the CTD oxygen and bottle dissolved oxygen in the latter part of Leg 2. Red dots in upper two panels indicate the result of calibration. Lower two panels show histogram of the difference between calibrated oxygen and bottle oxygen.

References

McTaggart, K. E., G. C. Johnson, M.C.Johnson, F.M.Delahoyde, and J.H.Swift (2010): The GO-SHIP Repeat Hydrography Manual: A Collection of Expert Reports and guidelines. IOCCP Report No 14, ICPO Publication Series No. 134, version 1, 2010

Budeus. G., and W. Schneider (1998): In-situ temperature calibration: A remark on instruments and methods. International WOCE Newsletter, No.30, WOCE International Project Office, Southampton, United Kingdom, 16-18.

Larson, N., and A.M. Pedersen (1996): Temperature measurements in flowing water: Viscous heating of sensor tips. Proc. of the First IGHEM Meeting, Montreal, QC, Canada, International Group for Hydraulic Efficiency Measurement. [Available online at http://www.seabird.com/technical_references/viscous.htm]

Uchida, H., K. Ohyama, S. Ozawa, and M. Fukasawa (2007): In-situ calibration of the Sea-Bird 9plus CTD thermometer. J. Atmos. Oceanic Technol.24, 1961-1967.

Akaike, H. (1974): A new look at the statistical model identification. IEEE Transactions on Automatic Control, 19:716 – 722.

Uchida, H., T. Kawano, I. Kaneko, and M. Fukasawa (2008): In –situ calibration of optode-based oxygen sensors. J. Atmos. Oceanic Technol., 25, 2271-2281.

Garcìa, H. E., and L. I. Gordon (1992): Oxygen solubility in seawater: Better fitting equations. Limnol. Oceanogr., 37, 1307-1312.

Shanno, David F. (1970): Conditioning of quasi-Newton methods for function minimization. Math. Comput. 24, 647-656. MR 42 #8905

Sea-Bird Electronics (2009): SBE 43 dissolved oxygen (DO) sensor – hysteresis corrections, Application note no. 64-3, 7 pp.

- Bottle Salinity

13 December 2013

(1)  Personnel

Keizo SHUTTA (GEMD/JMA)

Takahiro SEGAWA (GEMD/JMA)

Yasuaki BUNGI (GEMD/JMA)

Tetsuya NAKAMURA (GEMD/JMA)

Yoshikazu HIGASHI (GEMD/JMA)

Tomoyuki KITAMURA (GEMD/JMA)

(2)  Station occupied

A total of 68 stations (Leg 1: 40, Leg 2: 28) were occupied for bottle salinity. Station location and sampling layers of bottle salinity are shown in Figure C.2.1.

Figure C.2.1. Station location (left panel) and sampling layers of bottle salinity (right panel).

(3)  Instruments and method

(3.1) Salinity sample collection

The bottles in which the salinity samples are collected and stored are 250 ml colorless transparent glass bottles with screw caps. Each bottle was rinsed three times with sample water and was filled to the shoulder of the bottle. The screw caps were also thoroughly rinsed. Salinity samples were wiped with dry clothes and stored for more than 24 hours in the same laboratory as the salinity measurement was made.

(3.2) Instruments and methods

The salinity analysis was carried out on AUTOSAL Laboratory Salinometer model 8400B (S/N69677; Guildline Instruments Ltd., Canada), which was modified by addition of an Ocean Science International peristaltic-type sample intake pump and two Guildline platinum thermometers model 9450. One thermometer monitored an ambient temperature and the other monitored a bath temperature. The resolution of the thermometers was 1 mK. The measurement system was almost same as Aoyama et al (2003). The salinometer was operated in a ship's laboratory air-conditioned at a bath temperature of 24 deg-C. Ambient temperature varied from approximately 21.5 to 23.5 deg-C, while bath temperature is very stable and varied within ± 0.001 deg-C. A measure of a double conductivity ratio of a sample is taken as a median of thirty-one readings. Data collection was started after 10 seconds and it took about 10 seconds to collect 31 readings by a personal computer. Data were sampled for the fourth and the fifth filling of the cell. In case the difference in the double conductivity ratio between this two fillings was smaller than 0.00003, the average value of the two double conductivity ratios was used to calculate the bottle salinity with the algorithm for the Practical Salinity Scale, 1978 (hereafter PSS-78; UNESCO, 1981). If the difference was greater than or equal to 0.00003, we measured the sixth filling of the cell. In case the double conductivity ratio of the sixth filling did not satisfy the criteria above, we measured the next filling of the cell and chose proper two fillings which satisfied the criteria. We continued these process at most ninth fillings.

(4) Result

Standardization control was set to 5.72 and all the measurements were done by this setting. During the whole measurement, readings of STANDBY were almost 6019 ± 0001 and those of ZERO were 0.00001 or 0.00002. We used IAPSO Standard Seawater batch P152 whose conductivity ratio was 0.99981 (double conductivity ratio is 1.99962) as the standard for salinity. We measured 2 or 3 bottles of P152 for each station, total amount was 187. If some readings of SSW bottle were extremely high or low, we measured another bottle of SSW.

Figure C.2.2 shows the history of ambient temperature, bath temperature, double conductivity ratio of standard sea water (P152) and time drift of P152 readings but for four bad bottles. In raw P152 readings, it was found offset and long-term variability. To remove long-term variability, raw P152 readings were corrected to label value 1.99962 in the least significant digit of readings. After the correction, SSW drift was steady within 1 digit of readings in each leg. The average of corrected SSW double conductivity ratio was 1.999621 and the standard deviation was 0.000010, which was equivalent to 0.0002 in salinity. The same correction was applied to sample readings. The correction of AUTOSAL drift for salinity measurements was from 0 to 2 digits of readings.

During measurement of a sample taken at Stn.28 (RF3676), one heater lamp of the salinometer was broken down so the result of measurement was omitted from salinity calibration.

Figure C.2.2. The upper panel shows time-series of ambient temperature during cruise. The lower panel, black dots  and red dots indicate raw and corrected time-series of the double conductivity ratio of the standard sea water (P152), red line indicate linear regression of corrected standard sea water for each leg, gray line indicate label value double conductivity ratio of P152 and blue line indicates time-series of bath temperature during cruise.

(5) Sub-Standard Water

We also used sub-standard seawater which was filtered by pore size of 10 micrometer and stored in a 20 liter cubitainer made of polyethylene and stirred for at least 24 hours before measuring. It was measured every six samples in order to check possible sudden drift of the salinometer. During the whole measurements, there was no detectable sudden drift of the salinometer except for measurement of a sample taken at Stn.28 (RF3676).

(6) Replicate and Duplicate Samples

We took 267 pairs of replicate samples and 210 pairs of duplicate samples during the cruise. Figure C.2.3 and Figure C.2.4 show the absolute difference among replicate and duplicate samples in salinity, respectively. There were 30 bad measurements and 2 questionable measurements in replicate pairs and 23 bad measurements, 4 questionable measurements and 1 failure of sampling in duplicate pairs. Excluding those bad and questionable measurements, the mean absolute difference and standard deviation in 237 pairs of replicate samples was 0.0004±0.0004 in salinity and that in 182 pairs of duplicate samples was 0.0005±0.0006 in salinity. Note that standard deviation was calculated by a procedure (SOP23) in DOE (1994).

Table C.2.1. Summary of assigned quality control flags

| Flag | Definition | Salinity |
| --- | --- | --- |
| 2 | Good | 1716 |
| 3 | Questionable | 34 |
| 4 | Bad (Faulty) | 290 |
| 6 | Replicate measurements | 237 |
| Total number of samples |  | 2277 |

Figure C.2.3. Result of replicate samplings during this cruise against (a) station number, (b) sampling pressure and (c) salinity. Dotted line denotes the average of replicate samplings. Bottom panel (d) shows histogram of the result of replicate samplings.

Figure C.2.4. Same as Fig.C.2.3 but for duplicate samplings.

References

Aoyama, M., T. Joyce, T. Kawano and Y. Takatsuki (2003): Standards seawater comparison up to P129. Deep-sea Research, 1, Vol. 49, 1103-1114.

UNESCO (1981): Tenth report of the Joint Panel on Oceanographic Tables and Standards. UNESCO Tech. Papers in Mar. Sci., 36, 25 pp.

DOE (1994): Handbook of methods for the analysis of the various parameters of the carbon dioxide system in sea water; version 2. A.G. Dickson and C. Goyet (eds), ORNL/CDIAC-74.

- Bottle Oxygen

13 December 2013

(1) Personnel

Yusuke TAKATANI (GEMD/JMA)

Shinichiro UMEDA (GEMD/JMA)

Sho HIBINO (GEMD/JMA)

(2) Station occupied

A total of 71 stations (Leg 1: 42, Leg 2: 29) were occupied for bottle oxygen. Station location and sampling layers of bottle oxygen are shown in Figure C.3.1.

Figure C.3.1. Station location (left panel) and sampling layers of bottle oxygen (right panels).

(3) Reagents

· Manganous chloride solution (3 M) (Pickling Reagent-I)

Dissolved 600 g of MnCl2·4H2O in deionized water, then dilute the solution with deionized water to a final volume of 1 L. MnCl2·4H2O (Lot. CDP6460) used to make pickling reagent-I was guaranteed reagent manufactured by Wako Pure Chemical industries, Ltd.

· Sodium hydroxide (8 M) / sodium iodide solution (4 M) (Pickling Reagent-II)

Dissolved 320 g of NaOH in about 500 ml of deionized water, allow to cool, then add 600 g NaI and dilute with deionized water to a final volume of 1 L. NaOH (Lot. STN1103) and NaI (Lot. STQ5226) used to make pickling reagent-II were guaranteed reagent manufactured by Wako Pure Chemical industries, Ltd.

· Sulfuric acid solution (5 M)

Slowly add 280 ml concentrated H2SO4 to roughly 500 ml of deionized water. After cooling the final volume should be 1 L. H2SO4 (Lot. KWK1803) used to make sulfuric acid solution was guaranteed reagent manufactured by Wako Pure Chemical industries, Ltd.

· Sodium thiosulfate (0.04 M)

Dissolved 50 g of Na2S2O3·5H2O and 0.4 g of Na2CO3 in deionized water, then dilute the solution with deionized water to a final volume of 5 L. Na2S2O3·5H2O (Lot. PER3227) and Na2CO3 (Lot.WKF1312) used to make sodium thiosulfate were guaranteed reagent manufactured by Wako Pure Chemical industries, Ltd.

·Potassium iodate (0.001667 M)

Dry high purity KIO3 for two hours in an oven at 130°C. After weight out accurately KIO3, dissolve it in deionized water in a 5 L flask. Concentration of potassium iodate is determined by a gravimetric method. KIO3 (Lot. 62404E) used to make potassium iodate was manufactured by MERCK & CO., Inc., and a purity of KIO3 that is traceable to NIST (National Institute of Standards and Technology) standard reference material is 99.75±0.05 %. The normality of the standard potassium iodate solution made by Merck reagent was corrected by the factor as 1.0026. Details about this correction are described in chapter C.3 (16).

(4) Instruments

Detector;

DOT-01X automatic photometric titrator manufactured by Kimoto Electronic Co. Ltd.

Burette for sodium thiosulfate;

APB-510 manufactured by Kyoto Electronic Co. Ltd. / 10 ml of titration vessel

Burette for potassium iodate;

Multipette stream 4986 and Combitip plus manufactured by eppendorf /10 ml of tip vessel

Bottle top dispenser for pickling reagent-I and II;

CalibrexTM 520 manufactured by SOCOREX ISBA S.A.

(5) Seawater sampling

Following procedure is based on a determination method in IOCCP Report No.14 (Langdon, 2010). Seawater samples were collected from 10-liters Niskin bottles attached the CTD-system and a stainless steel bucket for the surface. Seawater for bottle oxygen measurement was transferred from the Niskin sample bottle and a stainless steel bucket to a volumetrically calibrated dry glass bottles (ca. 120 ml, standard deviation of calibration = 0.008 ml). At least three times volume of the glass of sample water was overflowed. Two reagent solutions (Reagent-I and II) of 1 ml (standard deviation of calibration = 0.003 ml) each were added immediately, sample temperature was then measured by a thermometer. After the stopper was inserted carefully into the glass, the sample glass was shaken vigorously to mix the content and to disperse the precipitate finely. The precipitate has settled at least halfway down the glass, the glass was then shaken again vigorously to disperse the precipitate. The sample glasses containing pickled samples were stored in a laboratory until they were titrated. To prevent air from entering the flask, deionized water was added to the neck of the flask after sampling.

(6) Sample measurement

At least 30 minutes after the re-shaking, the pickled samples were measured on board. 1 ml sulfuric acid solution and a magnetic stirrer bar were added into the sample glass and stirring began. Samples were titrated by sodium thiosulfate solution whose molarity was determined by potassium iodate solution. Temperature of sodium thiosulfate during titration was recorded by a thermometer. The titrations were carried out using the titration apparatus, named DOT-01X. Dissolved oxygen concentration (mol/kg) was calculated by the sample temperature at the fixation, CTD salinity, glass volume, and titrated volume of the sodium thiosulfate solution.

(7) Standardization

Concentration of sodium thiosulfate titrant (ca. 0.04 M) was determined by potassium iodate solution. Table C.3.1 shows a list of potassium iodate solution used in this cruise. Using a calibrated volumetric dispenser, 10 ml (standard deviation of calibration = 0.006 ml) of the standard potassium iodate solution was added to a glass with 100 ml of deionized water. Then, 1 ml of sulfuric acid solution, and 1 ml of pickling reagent solution-II and I were added into the glass in order. Amount of titrated volume of sodium thiosulfate (usually 5 times measurements average) gave the molarity of the sodium thiosulfate titrant. Figure C.3.2 and Table C.3.2 show the results of the standardization during this cruise. The sodium thiosulfate titrant of each batch was a mean of titrated volume of sodium thiosulfate on each day and a standard deviation of a concentration at 20 °C of sodium thiosulfate on each day was an uncertainty caused by the standardization. A sodium thiosulfate of one batch was assumed to be one sodium thiosulfate titrant. The uncertainty of dissolved oxygen that caused by the standardization was estimated 0.01-0.13 %.

Table C.3.1. List of the standard potassium iodate solution in this cruise.

| KIO3 batch | Conc. at 20°C (N) |
| --- | --- |
| KIO3_I | 0.009980±0.000003 |
| KIO3_IV | 0.009993±0.000003 |

Figure C.3.2. Results of the standardization. Upper panel shows results of end point, bottom panel shows results of calculated concentration at 20°C of sodium thiosulfate. Crosses show each value for each standardization samples, and closed circles show the mean at each standardizations. Thick lines and dotted lines denote the means and 1  error for each batch of sodium thiosulfate, respectively.

Table C.3.2. Results of the standardization.

| Leg | Date (UTC) | KIO3 |  | Na2S2O3 (ml) |  | Stations |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Batch | Bottle | Batch | End Point |  |
| 1 | 2010/7/6 | KIO3_ IV | 2 | #1 | 2.4359 | Stn.1 |
|  | 2010/7/11 | KIO3_ IV | 3 | #1 | 2.4423 | ¦ |
|  | 2010/7/15 | KIO3_ IV | 4 | #1 | 2.4372 | Stn.43 |
|  | 2010/7/17 | KIO3_ IV | 5 | #1 | 2.4384 |  |
|  |  | Na2S2O3_#1 |  | 2.4385±0.0028 |  |  |
|  | 2010/7/17 | KIO3_ IV | 5 | #2 | 2.4392 | Stn.44 |
|  | 2010/7/21 | KIO3_ IV | 6 | #2 | 2.4371 | ¦ |
|  | 2010/7/24 | KIO3_ IV | 7 | #2 | 2.4390 | Stn.67 |
|  |  | Na2S2O3_#2_1 |  | 2.4384±0.0012 |  |  |
| 2 | 2010/8/1 | KIO3_ IV | 9 | #2 | 2.4431 | Stn.68- |
|  | 2010/8/5 | KIO3_ IV | 10 | #2 | 2.4427 | Stn.84 |
|  |  | Na2S2O3_#2_2 |  | 2.4429±0.0003 |  |  |
|  | 2010/8/5 | KIO3_ IV | 10 | #3 | 2.4442 | Stn.93 |
|  | 2010/8/10 | KIO3_ IV | 11 | #3 | 2.4415 | ¦ |
|  | 2010/8/16 | KIO3_ IV | 12 | #3 | 2.4381 | Stn.105 |
|  |  | Na2S2O3_#4 |  | 2.4413±0.0031 |  |  |

(8) Determination of the blank

The oxygen in the pickling reagents-I (1 ml) and II (1 ml) was assumed to be 7.6 x 10-8 mol (Murray et al., 1968). The blank from the presence of redox species apart from oxygen in the reagents (the pickling reagents-I, II, and the sulfuric acid solution) was determined as follows. Using a calibrated volumetric dispenser, 1 ml of the standard potassium iodate solution was added to a glass with 100 ml of deionized water. Then, 1 ml of sulfuric acid solution, and 1 ml of pickling reagent solution-II and I were added into the glass in order. First, the sample was titrated to the end-point by sodium thiosulfate solution. Then, the sample was titrated again to the end-point after added a further 1 ml of the standard potassium iodate solution. The blank was determined by difference between the first (1 ml of KIO3) titrated volume of the sodium thiosulfate and the second (2 ml of KIO3) one. Because reagents set were prepared two sets (set A and B), the blank in each sets were determined. Usually, the results of 5 times blank determinations were averaged (Table C.3.3). The standard deviation of the blank determination during this cruise was 0.0010 (set A) and 0.0011 (set B) ml, c.a. 0.02 %.

Table C.3.3. Result of the blank determinations.

| Date (UTC) | Na2S2O3 | Blank (ml) |  | Samples (stations) |
| --- | --- | --- | --- | --- |
|  | Batch | Set A | Set B |  |
| 2010/7/6 | #1 | 0.0008 | 0.0004 | Stn.1-Stn.12 |
| 2010/7/9 | #1 | 0.0014 | 0.0029 | Stn.13-Stn.18 |
| 2010/7/11 | #1 | 0.0024 | 0.0029 | Stn.19-Stn.26 |
| 2010/7/13 | #1 | 0.0018 | 0.0013 | Stn.27-Stn.32 |
| 2010/7/15 | #1 | 0.0021 | 0.0023 | Stn.33-Stn.42 |
| 2010/7/17 | #2 | 0.0015 | 0.0020 | Stn.43-Stn.57 |
| 2010/7/21 | #2 | 0.0015 | 0.0015 | Stn.58-Stn.67 |
| 2010/8/1 | #2 | 0.0042 | 0.0042 | Stn.68-Stn.83 |
| 2010/8/5 | #3 | 0.0028 | 0.0031 | Stn.84-Stn.103 |
| 2010/8/10 | #3 | 0.0027 | 0.0033 | Stn.104-Stn.124 |

(9) Reagent blank

The blank determined in section (8), pure water blank (Vblk, dw) can be represented by equation (i),

Vblk, dw = Vblk, ep + Vblk, reg                (i)

where

Vblk, ep = blank due to differences between the measured end-point and the equivalence point;

Vblk, reg = blank due to oxidants or reductants in the reagent.

Here, the reagent blank (Vblk, reg) was determined by following procedure. 1 ml of the standard potassium iodate solution and 100 ml of deionized water were added to two glasses each. 1 ml of sulfuric acid solution, pickling reagent solution-II and I each were added into the first glass in order. Then, two times volume of the reagents (2 ml of sulfuric acid solution, pickling reagent solution-II and I each) was added to the second glass. The reagent blank was determined by difference between the first (3 ml of the total reagent volume added) titrated volume of the sodium thiosulfate and the second (6 ml of the total reagent volume added) one. We also carried out experiments for three and four times volume of the reagents. The results are shown in Figure C.3.3.

The relation between difference of the titrant (Na2S2O3) volume and the volume of the reagents added (Vreagent) is expressed by equation (ii),

Difference of the titrant volume = –0.0023Vreagent      (ii)

Vblk, reg was estimated to be about –0.007 ml, suggesting that about 0.04 mol of reductants was contained in every 3 ml of the reagents added.

Figure C.3.3. Blank (ml) due to redox species apart from oxygen in the reagents.

(10) Sample blank

Blank due to redox species other than oxygen in the sample (Vblk, spl) can be a potential source of measurement error. The total blank during the seawater measurement, the seawater blank (Vblk, sw) can be represented by equation (iii).

Vblk, sw = Vblk, spl + Vblk, dw         (iii)

If the pure water blank (Vblk, dw) that is determined in section (9) is identical both in pure water and in seawater, the difference between the seawater blank and the pure water one gives the sample blank (Vblk, spl).

Here, Vblk, spl was determined by following procedure. Seawater sample was collected in the calibrated volumetric glass (c.a. 120 ml) without the pickling. Then 1 ml of the standard potassium iodate solution, sulfuric acid solution, and pickling reagent solution-II and I each were added into the glass in order. Additionally, a glass contained 100 ml of deionized water and 1 ml of the standard potassium iodate solution, sulfuric acid solution, pickling reagent solution-II and I was prepared. The difference of the titrant volumes of the seawater glass and the deionized water one gave the sample blank (Vblk, spl).

We measured vertical profiles of the sample blank at 3 stations (Table C.3.4). The sample blank ranged from 0.17 to 1.96 mol/kg and its vertical and horizontal variations are large. This result does not agree to reported values ranged from 0.4 to 0.8 mol/kg (Culberson et al., 1991). It does not have been known about the magnitude and variability of the seawater blank, so this result should be discussed carefully. Ignorant of the sample blank will cause systematic errors in the oxygen calculations, but these errors are expected to be the same to all investigators and not to affect the comparison of results from different investigators (Culberson, 1994).

Table C.3.4. Results of the sample blank determinations during this cruise.

| Station: Stn.75 4.45˚N/139.53˚E |  | Station: Stn.111 2.50˚N/139.50˚E |  | Station: Stn.115 2.91˚N/137.50˚E |  |
| --- | --- | --- | --- | --- | --- |
| Pres. (dbar) | Blank (mol/kg) | Pres. (dbar) | Blank (mol/kg) | Pres. (dbar) | Blank (mol/kg) |
| 10.3 | 0.17 | 10.3 | 0.35 | 25.3 | 0.91 |
| 50.6 | 1.03 | 103.5 | 0.74 | 100.8 | 0.63 |
| 101.7 | 0.54 | 201.2 | 1.04 | 302.0 | 0.95 |
| 251.5 | – | 381.9 | 0.70 | 503.2 | 1.09 |
| 634.7 | 0.61 | 483.8 | 1.76 | 1,008.8 | 0.98 |
| 838.0 | 0.90 | 675.1 | 1.96 | 1,514.0 | 0.75 |
| 1,596.4 | 0.56 | 978.3 | 0.98 | 2,020.6 | 1.29 |
| 2,611.7 | 0.82 | 1,940.0 | 1.39 | 3,038.5 | 1.44 |
| 3,631.6 | 0.71 | 2,955.6 | 1.11 | 4,060.8 | 1.29 |
| 4,142.7 | 0.84 | 4,382.1 | – | 4,574.2 | 0.90 |

(11) Replicate sample measurement

Replicate samples were carried out at every bottle oxygen observation stations. Total amount of the replicate sample pairs in good measurement (flag=6) was 234, and total amount of the removed pair (flag=3 or 4) was 14. The average and the standard deviation of the replicate measurement during this cruise were 0.17±0.17 mol/kg. The standard deviation was calculated by a procedure (SOP23) in DOE (1994). The difference between the replicate sample pairs did not depended on sampling pressure, measurement date and concentration of sample (Figure C.3.4). The averages and the standard deviations during Leg 1 and Leg 2 were 0.17±0.17 (n=127) and 0.18±0.16 (n=107) mol/kg, respectively.

Figure C.3.4. Result of replicate samplings during this cruise against (a) station number, (b) sampling pressure and (c) concentration of dissolved oxygen. Dotted line denotes the average of replicate samplings. Bottom panel (d) shows histogram of the result of replicate samplings.

(12) Duplicate sample measurement

Duplicate samples that were seawater samples from two Niskin sample bottles that were collected at same depth were carried out at almost every bottle oxygen observation stations also. Total amount of the duplicate sample pairs in good measurement (flag=2) was 156, and total amount of the removed pair (flag=3 or 4) was 16. The average and the standard deviation of the duplicate measurement during this cruise were 0.22±0.21 mol/kg. The difference between the duplicate sample pairs did not depended on measurement date and concentration of sample, but the results of the duplicate measurement on surface and subsurface (above 1,000 dbar) were large (Figure C.3.5). We thought that this reason was because water mass on a surface and a subsurface was not similar compared with that of an intermediate and a deep layer. The averages and the standard deviations during Leg 1 and Leg 2 were 0.26±0.24 (n=60) and 0.20±0.18 (n=96) mol/kg, respectively.

Figure C.3.5. Result of duplicate samplings during this cruise against (a) station number, (b) sampling pressure and (c) concentration of dissolved oxygen. Dotted line denotes the average of duplicate samplings. Bottom panel (d) shows histogram of the result of duplicate samplings.

(13) Mutual comparison between each standard potassium iodate

During the cruise, we performed the mutual comparison between two standard potassium iodate of difference Lot. in order to confirm the accuracy of our oxygen measurement and the bias of a standard potassium iodate. We measured concentration of a KIO3 (KIO3_I) against another KIO3 (KIO3_IV), and checked the difference between measurement value and theoretical one (Table C.3.5, Figure C.3.6). Error weighted means of measurement results of KIO3_I were 0.009947±0.000012 N. The averaged value of the KIO3_I was so close to the theoretical value (0.009955±0.000003 N) that was prepared in laboratory. A good agreement among two standard potassium iodate confirmed that there was no systematic shift in our oxygen measurements during this cruise.

Table C.3.5. Results of mutual comparison of KIO3_I against KIO3_IV

| Date (UTC) | KIO3 Batch | Measurement Value (N) |
| --- | --- | --- |
| 2010/7/6 | KIO3_I_4 | 0.009982±0.000013 |
| 2010/7/11 | KIO3_I_5 | 0.009966±0.000011 |
| 2010/7/15 | KIO3_I_6 | 0.009970±0.000011 |
| 2010/7/17 | KIO3_I_7 | 0.009968±0.000012 |
| 2010/7/17 | KIO3_I_7 | 0.009964±0.000012 |
| 2010/7/21 | KIO3_I_8 | 0.009977±0.000012 |
| 2010/7/24 | KIO3_I_9 | 0.009978±0.000013 |
| 2010/8/1 | KIO3_I_10 | 0.009981±0.000011 |
| 2010/8/5 | KIO3_I_11 | 0.009968±0.000011 |
| 2010/8/5 | KIO3_I_11 | 0.009966±0.000013 |
| 2010/8/10 | KIO3_I_12 | 0.009983±0.000012 |
| 2010/8/16 | KIO3_I_13 | 0.009982±0.000012 |
| Weighted Mean |  | 0.009973±0.000012 |

Figure C.3.6. Results of mutual comparison of KIO3_I against KIO3_IV. Closed circles show mean of measurement value with 1  error at each mutual comparison, and gray opened diamonds and error bar show the theoretical value and the uncertainty of the standard potassium iodate.

(14) Quality control flag assignment

Quality flag values were assigned to oxygen measurements using the code defined in IOCCP Report No.14 (Swift, 2010). Measurement flags of 2 (good), 3 (questionable), 4 (bad), 5 (not reported), and 6 (replicate measurement) have been assigned (Table C.3.6). The replicate data were averaged and flagged 6 if both of them were flagged 2. If either of them was flagged 3 or 4, a datum with “younger” flag was selected. For the choice between 2, 3, or 4, we basically followed a flagging procedure as listed below:

- Vertical sections against pressure and potential density were drawn. Any points not lying on a generally smooth curve were noted.

- Dissolved oxygen was then plotted against potential temperature, salinity and nutrients. If a datum deviated from a group of plots, it was flagged 3.

- If a datum was deviated from the mean ±3 on the section, datum flag was degraded from 2 to 3, or from 3 to 4.

- We Compared bottle oxygen with CTD oxygen at the sampling depth. If a datum deviated from CTD oxygen, datum flag was degraded from 2 to 3, or from 3 to 4.

- If the bottle flag was 4 (did not trip correctly), a datum was flagged 4 (bad). If the bottle flag was 3 (leaking) or 5 (unknown problem), a datum was flagged based on steps a, b, c, and d.

Table C.3.6. Summary of assigned quality control flags.

| Flag | Definition | Number |
| --- | --- | --- |
| 2 | Good | 2353 |
| 3 | Questionable | 326 |
| 4 | Bad (Faulty) | 52 |
| 5 | Not reported | 1 |
| 6 | Replicate measurements | 234* |
| Total number of samples |  | 2732* |

*Samples of flag 6 are counted as flag 2

(15) Problems

A leak of sodium thiosulfate from a joint of burette was revealed at the standardization after Stn.67. From the comparison between bottle oxygen and the oxygen sensor (RINKO III manufactured by JFE Advantech Co. Ltd.), it was thought that this problem had happened from Stn.58 (Figure C.3.7). The data between Stn.58 and Stn.67 were flagged 3 or 4.

Figure C.3.7. Comparison between bottle oxygen and RINKO III. Closed circles show data between Stn.58 and Stn.67.

(16) Correction of a standard potassium iodate made by Merck reagent

We have conducted the inter-laboratory comparison of the standard potassium iodate solution between JMA and Japan Agency for Marine-Earth Science and Technology (JAMSTEC) in order to ensure the comparability of dissolved oxygen concentration in sea water since 2010. Moreover, we also conducted the inter-laboratory comparison of the standard potassium iodate solution between JMA and Scripps Institution of Oceanography (SIO) on board R/V Melville in May 2013. Table C.3.7 shows pure KIO3 reagents used in each laboratory.

The results of inter-laboratory comparisons are summarized as follows. The standard potassium iodate solution made by Merck reagent shows a systematic difference to the solutions made from Wako reagent and Alfa Aesar reagent. On the other hand, the standard potassium iodate solution made by National Metrology Institute of Japan (NMIJ) reagent has no significant difference to the standard solutions made from Wako reagent and Alfa Aesar reagent (Table C.3.8 and Figure C.3.8).

The measured normality of standard potassium iodate solution made by Merck reagent based on the nominal normality of NMIJ reagent is higher (0.26±0.12%) than the nominal normality (Figure C.3.9). From t-test, this difference is significant. Thus, the normality of the standard potassium iodate solution made by Merck reagent should be corrected by the factor as 1.0026 from the result of the inter-laboratory comparison studies conducted by JMA.

Table C.3.7. List of pure KIO3 reagent used in each institution

| Institution | Company | Lot. | Purity (%) |
| --- | --- | --- | --- |
| JMA | Merck | 62404E | 99.75 (2010~2011) |
|  |  | 92404G | 99.74 (2011~) |
|  | NMIJ | CRM 3006-a | 99.973 (2012~) |
| JAMSTEC | Wako | EPR3227 | >99.95 |
| SIO | Alfa Aesar | B05N35 | 99.4 ~ 100.4 |

Table C.3.8 The results of standard potassium iodate solution inter-comparison.

| Date | measured (N) | Uncertainty (N) | Nominal (N) | Uncertainty*** (N) | Ratio (measured/nominal) |
| --- | --- | --- | --- | --- | --- |
| (1) normality of the standard potassium iodate solution made by Wako versus Merck at JMA |  |  |  |  |  |
| 2010/5/23 | 0.009985 | 0.000008* | － | － | － |
| 2010/5/29 | 0.010004 | 0.000007* | － | － | － |
| 2010/12/18 | 0.009983 | 0.000009* | － | － | － |
| 2011/6/22 | 0.009981 | 0.000013* | 0.010014 | － | 0.9967 |
| 2011/6/22 | 0.009978 | 0.000013* | 0.010014 | － | 0.9964 |
| 2011/6/22 | 0.009986 | 0.000012* | 0.010014 | － | 0.9972 |
| 2011/6/22 | 0.009981 | 0.000012* | 0.010014 | － | 0.9967 |
| 2011/6/22 | 0.009984 | 0.000013* | 0.010014 | － | 0.9970 |
| 2011/6/22 | 0.009988 | 0.000012* | 0.010014 | － | 0.9974 |
| 2012/5/28 | 0.009993 | 0.000008* | 0.010009 | － | 0.9984 |
| 2012/5/28 | 0.009986 | 0.000009* | 0.010010 | － | 0.9976 |
| 2012/5/28 | 0.009989 | 0.000008* | 0.010012 | － | 0.9977 |
| 2012/5/28 | 0.009986 | 0.000009* | 0.010012 | － | 0.9974 |
| 2012/5/28 | 0.009989 | 0.000008* | 0.010012 | － | 0.9977 |
| 2012/5/28 | 0.009990 | 0.000008* | 0.010012 | － | 0.9978 |
| 2013/6/10 | 0.009976 | 0.000009* | 0.010002 | － | 0.9974 |
| 2013/6/10 | 0.009979 | 0.000009* | 0.010002 | － | 0.9977 |
| 2013/6/10 | 0.009974 | 0.000009* | 0.010002 | － | 0.9972 |
| Mean ratio of the measured normality versus the nominal normality = 0.9974 ± 0.0005 |  |  |  |  |  |
| (2) normality of the standard potassium iodate solution made by Wako versus NMIJ at JMA |  |  |  |  |  |
| 2012/5/28 | 0.010015 | 0.000008* | 0.010009 | － | 1.0006 |
| 2012/5/28 | 0.010008 | 0.000009* | 0.010010 | － | 0.9998 |
| 2012/5/28 | 0.010012 | 0.000008* | 0.010012 | － | 1.0000 |
| 2012/5/28 | 0.010008 | 0.000009* | 0.010012 | － | 0.9996 |
| 2012/5/28 | 0.010011 | 0.000008* | 0.010012 | － | 0.9999 |
| 2012/5/28 | 0.010013 | 0.000008* | 0.010012 | － | 1.0001 |
| 2013/6/10 | 0.010001 | 0.000009* | 0.010002 | － | 0.9999 |
| 2013/6/10 | 0.010003 | 0.000010* | 0.010002 | － | 1.0001 |
| 2013/6/10 | 0.009999 | 0.000009* | 0.010002 | － | 0.9997 |
| Mean ratio of the measured normality versus the nominal normality = 1.0000 ± 0.0003 |  |  |  |  |  |

*: The standard uncertainty estimated theoretically by precision of burette, etc.

**: The standard deviation of 4-5 times measurements

***: The standard uncertainty by gravimetric method

Table C.3.8 (continued)

| Date | measured (N) | Uncertainty (N) | Nominal (N) | Uncertainty*** (N) | Ratio (measured/nominal) |
| --- | --- | --- | --- | --- | --- |
| (3) normality of the standard potassium iodate solution made by Merck versus Wako at JAMSTEC |  |  |  |  |  |
| 2010/7/14 | 0.009976 | － | 0.009955 | 0.000003 | 1.0021 |
| 2010/7/14 | 0.009990 | － | 0.009972 | 0.000003 | 1.0018 |
| 2010/7/14 | 0.009990 | － | 0.009972 | 0.000003 | 1.0018 |
| 2010/7/14 | 0.010002 | － | 0.009977 | 0.000003 | 1.0025 |
| 2010/9/3 | 0.009994 | － | 0.009967 | 0.000003 | 1.0027 |
| 2010/9/3 | 0.009979 | － | 0.009952 | 0.000003 | 1.0027 |
| 2011/5/30 | 0.010250 | － | 0.010233 | 0.000003 | 1.0017 |
| 2011/5/30 | 0.010168 | － | 0.010154 | 0.000003 | 1.0014 |
| 2011/5/30 | 0.010017 | － | 0.010001 | 0.000003 | 1.0016 |
| 2011/5/31 | 0.010573 | － | 0.010546 | 0.000003 | 1.0026 |
| 2011/5/31 | 0.010571 | － | 0.010546 | 0.000003 | 1.0024 |
| 2011/5/31 | 0.010856 | － | 0.010840 | 0.000003 | 1.0015 |
| 2012/4/24 | 0.010168 | 0.000005** | 0.010148 | 0.000003 | 1.0020 |
| 2012/4/24 | 0.010049 | 0.000004** | 0.010031 | 0.000003 | 1.0018 |
| 2012/4/24 | 0.010050 | 0.000004** | 0.010031 | 0.000003 | 1.0019 |
| 2012/4/25 | 0.010112 | 0.000004** | 0.010094 | 0.000003 | 1.0018 |
| 2012/4/25 | 0.010189 | 0.000005** | 0.010162 | 0.000003 | 1.0027 |
| 2012/4/25 | 0.010018 | 0.000004** | 0.010001 | 0.000003 | 1.0017 |
| 2013/5/21 | 0.010010 | 0.000003** | 0.009989 | 0.000003 | 1.0021 |
| 2013/5/21 | 0.010012 | 0.000003** | 0.009989 | 0.000003 | 1.0023 |
| 2013/5/21 | 0.010006 | 0.000004** | 0.009989 | 0.000003 | 1.0017 |
| 2013/5/21 | 0.010014 | 0.000004** | 0.009990 | 0.000003 | 1.0024 |
| 2013/5/21 | 0.010011 | 0.000005** | 0.009990 | 0.000003 | 1.0021 |
| 2013/5/21 | 0.010002 | 0.000005** | 0.009990 | 0.000003 | 1.0012 |
| Mean ratio of the measured normality versus the nominal normality = 1.0020 ± 0.0004 |  |  |  |  |  |
| (4) normality of the standard potassium iodate solution made by NMIJ versus Wako at JAMSTEC |  |  |  |  |  |
| 2012/4/25 | 0.009997 | 0.000004** | 0.009997 | 0.000001 | 1.0000 |
| 2013/5/19 | 0.009995 | 0.000004** | 0.009994 | 0.000001 | 1.0001 |
| 2013/5/19 | 0.009994 | 0.000006** | 0.009994 | 0.000001 | 1.0000 |
| 2013/5/19 | 0.009980 | 0.000003** | 0.009994 | 0.000001 | 0.9986 |
| 2013/5/21 | 0.009981 | 0.000003** | 0.009994 | 0.000001 | 0.9987 |
| Mean ratio of the measured normality versus the nominal normality = 0.9995 ± 0.0008 |  |  |  |  |  |

*: The standard uncertainty estimated theoretically by precision of burette, etc.

**: The standard deviation of 4-5 times measurements

***: The standard uncertainty by gravimetric method

Table C.3.8 (continued)

| Date | measured (N) | Uncertainty (N) | Nominal (N) | Uncertainty*** (N) | Ratio (measured/nominal) |
| --- | --- | --- | --- | --- | --- |
| (5) normality of the standard potassium iodate solution made by Merck versus Alfa Aesar at SIO |  |  |  |  |  |
| 2013/5/10 | 0.010023 | － | 0.009996 | 0.000003 | 1.0027 |
| 2013/5/10 | 0.010062 | － | 0.010033 | 0.000003 | 1.0029 |
| 2013/5/10 | 0.010050 | － | 0.010033 | 0.000003 | 1.0017 |
| Mean ratio of the measured normality versus the nominal normality = 1.0024 ± 0.0006 |  |  |  |  |  |
| (6) normality of the standard potassium iodate solution made by NMIJ versus Alfa Aesar at SIO |  |  |  |  |  |
| 2013/5/10 | 0.009994 | － | 0.009994 | 0.000001 | 1.0000 |
| 2013/5/10 | 0.009990 | － | 0.009994 | 0.000001 | 0.9996 |
| Mean ratio of the measured normality versus the nominal normality = 0.9998 ± 0.0003 |  |  |  |  |  |

*: The standard uncertainty estimated theoretically by precision of burette, etc.

**: The standard deviation of 4-5 times measurements

***: The standard uncertainty by gravimetric method

Figure C.3.8 Results of standard potassium iodate solution inter-laboratory comparison among JMA, JAMSTEC and SIO. Black (Gray) closed circles denote measured/nominal ratio of Wako (Merck) versus Merck (Wako), and black (gray) closed diamonds denote measured/nominal ratio of Wako (NMIJ) versus NMIJ (Wako). Gray open circles (diamonds) denote measured/nominal ratio of Merck (NMIJ) versus Alfa Aesar.

Figure C.3.9. Results of inter-comparison between Merck and NMIJ. Upper panel shows measured/nominal ratio of Merck versus NMIJ for each standard potassium iodate solution batch serial, and lower panel shows the histogram of its ratios.

(17) Uncertainty in Oxygen data of this cruise

The reproducibility in this cruise determined by replicate samples and duplicate samples in section (11) and (12) was 0.17±0.17 mol/kg and 0.22±0.21 mol/kg. Bottle oxygen data in this cruise were calculated based on IOCCP Report No.14 (Langdon, 2010). In these results, various uncertainties were included (ex. standardization, calibration of glass bottles, precision of burette etc.). Considering these uncertainty that can be estimated theoretically, it was estimated that the standard uncertainty of bottle oxygen data in this cruise is about 0.43 mol/kg. However, it is impossible to estimate an accurate uncertainty because there is no reference material.

(18) Results

(18.1) Comparison at cross-stations during this cruise

Cross-stations during this cruise were two stations. The one was located at 2°N/142°E, the another was located 7°N/137°E.

At stations of Stn.83 (RF3731) and Stn.104 (RF3755), hydrocast sampling for dissolved oxygen was conducted two times at interval of about five days. Dissolved oxygen profiles of the two hydrocasts agreed well (Figure C.3.10). We compared interpolated data, because the sampling layers of the two hydrocasts were difference. In the layers deeper than 2000 dbar, difference of interpolated data between the two hydrocasts was calculated to be about 1.23±0.30 mol/kg. In these layers, difference of the oxygen sensor between the two hydrocasts was also about 1.1±0.6 mol/kg.

At stations of Stn.67 (RF3715), Stn.68 (RF3716), and Stn.124 (RF3772), hydrocast sampling for dissolved oxygen was conducted three times. Interval between the first and the second was about a week, interval between the second and the third was about two weeks. Dissolved oxygen profiles between the second and third hydrocasts agreed well, but the data of the first hydrocast had slightly larger than that of second and third hydrocast because of a leak of sodium thiosulfate from a joint of the burette (Figure C.3.7). In the layers deeper than 2000 dbar, difference between the second and third hydrocast was calculated to be about 0.11±0.64 mol/kg.

Figure C.3.10. Comparison of dissolved oxygen profiles between the first hydrocast (pluses), the second one (gray diamonds) and the third one (circles) at the cross-stations of (a) 2°N/142°E and (b) 7°N/137°E. Lines denote the profiles of the oxygen sensor.

(18.2) Comparison at cross-stations of WHP-P2 section in 1994 and 2004

We compared our oxygen data and one of WHP-P2 at a cross point, around 30°N/137°E. WHP-P2 line was observed two times, first in 1994 by R/V Shoyo belonged to Maritime Safety Agency of Japan (MSA) (present Japan Coast Guard (JCG)) and repeated in 2004 by R/V Melville belonged to Scripps Institution of Oceanography (SIO). Dissolved oxygen profiles between one in this cruise and in 2004 agreed well (Figure C.3.11), and the difference blow 2000 dbar is 0.70±0.74 mol/kg. But it was found that oxygen data in this cruise were significantly lower than those in 1994 in deep layers, the differences below 2000 dbar is –4.18±2.71mol/kg. This difference should be discussed carefully.

(18.3) Comparison at cross-stations of WHP-P3 section in 1985 and 2005/06

We compared our oxygen data and one of WHP-P3 at a cross point, around 24°N/137°E. WHP-P3 line was observed two times, first in 1985 by R/V Thomas G. Thompson belonged to SIO and repeated in 2005/06 by R/V Mirai belong to Japan Agency for Marine-Earth Science and Technology (JAMSTEC). Dissolved oxygen profiles between one in this cruise and in 2005/06 agreed well (Figure C.3.11). The differences between oxygen data below 2000 dbar in this cruise and these in 1985, or in 2005/06 are –1.39±0.56, 0.28±1.14 mol/kg, respectively. For the comparison with oxygen data in 1985, the offset value is larger than reported adjustments, about minus 0.3 mol/kg (Johnson et al., 2001; Gouretski and Jancke, 2001). For the comparison with oxygen data in 2005/06, it should also be noted that the relatively large difference in deep layer ranged from about 1,000 to 2000 dbar. Though it might be caused by the slight difference of the observation position, it is necessary to discussed it carefully.

Figure C.3.11. Comparison of dissolved oxygen profiles at cross-stations of (a) WHP-P2 and (b) WHP-P3. Pluses, gray diamonds and opened circles show the first observation, the second one and this cruise, respectively. Lines denote the profiles of the oxygen sensor.

(18.4) Comparison with WHP-P9 oxygen data in 1994

We compared oxygen data in this cruise and one of WHP-P9 in 1994. In deep layers in a wide region, dissolved oxygen have been decreased from 1994 (Figure C.3.12). Below 2000 m, the difference in average is calculated in –1.38±2.28 mol/kg (Figure C.3.13). This offset value is closed to reported adjustments, about –1 mol/kg for dissolved oxygen data of WHP-P9 (Johnson et al., 2001; Gouretski and Jancke, 2001). Therefore, it was thought that oxygen measurements in this cruise were conducted correctly.

Figure C.3.12. Difference of dissolved oxygen between 2010 and 1994 against water depth.

Figure C.3.13. Bottle oxygen data in 1994 (pluses) and 2010 (circles) below 1500 m (left panel) and difference of dissolved oxygen on the standard depth (right panel). Black closed circles denote mean of the differences with 1  error.

Acknowledgements

We are grateful to Michio Aoyama of the Meteorological Research Institute for helpful advice and discussion about correction of a standard potassium iodate. We also thank Yuichiro Kumamoto of JAMSTEC, Fuyuki Shibata of Marine Works Japan, James H. Swift and Susan M. Becker of SIO for providing their comparison results.

References

Cullberson, A.H. (1994), Dissolved oxygen, in WHPO Pub. 91-1 Rev. 1, November 1994, Woods Hole, Mass., USA.

Cullberson, A.H., G. Knapp, M.C. Stalcup, R.T. Williams, and F. Zemlyak (1991), A comparison of methods for the determination of dissolved oxygen in seawater, WHPO Pub. 91-2, August 1991, Woods Hole, Mass., USA.

DOE (1994), Handbook of methods for the analysis of the various parameters of the carbon dioxide system in sea water; version 2. A.G. Dickson and C. Goyet (eds), ORNL/CDIAC-74.

Gouretski, V.V. and K. Jancke (2001), Systematic errors as the causes for an apparent deep water property variability: global analysis of the WOCE and historical hydrographic data, Prog. Oceanogr., 48, 337-402.

JAMSTEC, WHP P03 REVISIT DATA BOOK (2001), edited by T. Kawano and H. Uchida, JAMSTEC

Johson, G.C., P.E. Robbins, and G.E. Hufford (2001), Systematic adjustments of hydrographic sections for internal consistency, J. Atomos. Oceanic Technol., 18, 1234-1244.

Murray, C.N., J.P. Riley and T.R.S. Wilson (1968), The solubility of oxygen in Winkler reagents used for the determination of dissolved oxygen, Deep-Sea Res., 15, 237-238

Langdon, C. (2010), Determination of dissolved oxygen in seawater by Winkler titration using the amperometric technique, IOCCP Report No.14, ICPO Pub. 134, 2010 ver.1

Swift, J. H. (2010): Reference-quality water sample data: Notes on acquisition, record keeping, and evaluation. IOCCP Report No.14, ICPO Pub. 134, 2010 ver.1

- Nutrients

Updated 8 July 2020

- Personnel

Kazuhiro SAITO (GEMD/JMA)

Hiroyuki FUJIWARA (GEMD/JMA)

Takahiro KITAGAWA (GEMD/JMA)

(2) Station occupied

A total of 104 stations (Leg1: 61, Leg2: 43) were occupied for nutrients. Station location and sampling layers of nutrients are shown in Figure C.4.1.

Figure C.4.1. Station location (left) and sampling layers (right) of nutrients.

(3) Instrument and Method

(3.1) Analytical detail using Auto Analyzer III systems (BLTEC)

The nutrients analyses were carried out on 4-channel Auto Analyzer III (BLTEC). Measured Parameters are nitrate + nitrite, nitrite, phosphate and silicate.

Nitrate + nitrite and nitrite are analyzed according to the modification method of Armstrong (1967). The sample nitrate is reduced to nitrite in a cadmium tube inside of which is coated with metallic copper. The sample stream with its equivalent nitrite is treated with an acidic, sulfanilamide reagent and the nitrite forms nitrous acid which reacts with the sulfanilamide to produce a diazonium ion. N-1-Naphthylethylene-diamine added to the sample stream then couples with the diazonium ion to produce a red, azo dye. With reduction of the nitrate to nitrite, both nitrate and nitrite react and are measured; without reduction, only nitrite reacts. Thus, for the nitrite analysis, no reduction is performed and the alkaline buffer is not necessary.

The phosphate analysis is a modification of the procedure of Murphy and Riley (1962). Molybdic acid is added to the seawater sample to form phosphomolybdic acid which is in turn reduced to phosphomolybdous acid using L-ascorbic acid as the reductant.

The silicate method is analogous to that described for phosphate. The method used is essentially that of Grasshoff et al. (1983), wherein silicomolybdic acid is first formed from the silicate in the sample and added molybdic acid, then the silicomolybdic acid is reduced to silicomolybdous acid, or "molybdenum blue," using L-ascorbic acid as the reductant.

The flow diagrams and reagents for each parameter are shown in Figures C.4.2-C.4.5.

(3.2) Nitrate Reagents

Ammonium chloride (buffer), 0.7 M (0.04 % w/v);

Dissolve 190 g Ammonium chloride, NH4Cl, in ca. 5000 ml of milli-Q water, add about 5 ml Ammonia (aq.), adjust pH 8.2-8.5.

Sulfanilamide, 0.06 M (1 % w/v);

Dissolve 5 g Sulfanilamide, 4-NH2C6H4SO3H, in 430 ml milli-Q water, add 70 ml concentrated HCl. After mixing, 1 ml Brij-35 (22 % w/w) is added.

N-1-Naphtylethylene-diamine dihydrochloride (NEDA), 0.004 M (0.1 % w/v);

Dissolve 0.5 g NEDA, C10H7NH2CH2CH2NH2·2HCl, in 500 ml milli-Q water.

Figure C.4.2. 1ch. (Nitrate + Nitrite) Flow diagram.

(3.3) Nitrite Reagents

Sulfanilamide, 0.06 M (1 % w/v);

Dissolve 5 g Sulfanilamide, 4-NH2C6H4SO3H, in 430 ml milli-Q water, add 70 ml concentrated HCl. After mixing, 1 ml Brij-35 (22 % w/w) is added.

N-1-Naphtylethylene-diamine dihydrochloride (NEDA), 0.004 M (0.1 % w/v);

Dissolve 0.5 g NEDA, C10H7NH2CH2CH2NH2·2HCl, in 500 ml milli-Q water.

Figure C.4.3. 2ch. Nitrite Flow diagram.

(3.4) Phosphate Reagents

Ammonium molybdate, 0.005 M (0.6 % w/v);

Dissolve 3 g Ammonium molybdate(VI) tetrahydrate, (NH4)6Mo7O24·4H2O, and 0.05 g Potassium antimonyl tartrate, C8H4K2O12Sb2·3H2O, in 400 ml milli-Q water and add 100 ml H2SO4 (12.6N). After mixing, 2 ml Sodium dodecyl sulfate (15 % solution in water) is added.

L(+)-Ascorbic acid, 0.08 M (1.5 % w/v);

Dissolve 4.5 g L(+)-Ascorbic acid, C6H8O6, in 300 ml milli-Q water. After mixing, 10 ml Acetone is added. Freshly prepared before every measurement.

Figure C.4.4. 3ch. Phosphate Flow diagram.

(3.5) Silicate Reagents

Ammonium molydate, 0.005 M (0.6 % w/v);

Dissolve 3 g Ammonium molybdate (VI) tetrahydrate, (NH4)6Mo7O24·4H2O, in 495 ml milli-Q water and added 5 ml H2SO4 (12.6N). After mixing, 2 ml Sodium dodecyl sulfate (15 % solution in water) is added.

Oxalic acid, 0.4 M (5 % w/v);

Dissolve 25 g Oxalic acid dihydrate, (COOH)2·2H2O, in 500 ml milli-Q water.

L(+)-Ascorbic acid, 0.08 M (1.5 % w/v);

Dissolve 4.5 g L(+)-Ascorbic acid, C6H8O6, in 300 ml milli-Q water. After mixing, 10 ml Acetone is added. Freshly prepared before every measurement.

Figure C.4.5. 4ch. Silicate Flow diagram.

(3.6) Sampling procedures

Seawater samples were collected from 10-liters Niskin bottle attached CTD-system and a stainless steel bucket for the surface. Sampling of nutrients followed that oxygen and trace gases. Samples were drawn into 10 ml polymethylpenten vials with sample drawing tubes. These were rinsed three times before filling and vials were capped immediately after the drawing.

No transfer was made and the vials were set an auto sampler tray directly. Samples were analyzed immediately after collection.

(3.7) Data processing

Raw data from Auto Analyzer III were recorded at 1-second interval and were treated as follows;

- Calculate 11-second moving average.

- Check the shape of each peak and position of peak values taken, and then change the positions of peak values taken if necessary.

- Baseline correction was done basically using liner regression.

- Reagent blank correction was done basically using liner regression.

- Carry-over correction was applied to peak heights of each sample.

- Sensitivity correction was applied to peak heights of each sample.

- Refraction error correction was applied to peak heights of each seawater sample.

- Calibration curves to get nutrients concentration were assumed quadratic expression.

- Load pressure and salinity from CTD data to calculate density of seawater.

- Convert data from μmol/l to μmol/kg.

(4) Nutrients standards

(4.1) Volumetric Laboratory Ware of in-house standards

All volumetric ware used were gravimetrically calibrated. Polymethylpenten volumetric flasks were gravimetrically calibrated at the temperature of use within 3 - 4 K.

Volumetric flasks

The weights obtained in the calibration weightings were corrected for the density of water and air buoyancy.

Pipettes and pipettors

All pipettes have nominal calibration tolerances of 0.1 % or better. These were gravimetrically calibrated in order to verify and improve upon this nominal tolerance.

(4.2) Reagents, general considerations

Specifications

For nitrate standard, “potassium nitrate 99.995 suprapur” provided by Merck, CAS No. : 7757-79-1, was used.

For phosphate standard, “potassium dihydrogen phosphate anhydrous 99.995 suprapur” provided by Merck, CAS No. : 7778-77-0, was used.

For nitrite standard, “sodium nitrite GR for analysis ACS,Reag. Ph Eur” provided by Merck, CAS No. : 7632-00-0, was used.

For the silicate standard, we use “Silicon standard solution traceable to SRM from NIST SiO2 in NaOH 0.5 mol/L 1000 mg/L Si CertiPUR” provided by Merck, which lot number is HC814662 is used. The silicate concentration is certified by NIST-SRM3150 as 999 mg/L with the expanded uncertainty of 5 mg/L (k=2). However, we used concentration corrected with factor as 0.975. Details about this correction are described in chapter C4 (10).

Ultra pure water

Ultra pure water (Milli-Q water) freshly drawn was used for preparation of reagents, higher concentration standards and for measurement of reagent and system blanks.

Low Nutrient Seawater (LNSW)

Surface water having low nutrient concentration was taken and filtered using 10 μm pore size membrane filter. This water is stored in 20 liter flexible container with paper box.

(4.3) Concentrations of nutrients for A, B and C standards

Concentrations of nutrients for A, B and C standards are set as shown in Table C.4.1. The C standard is prepared according recipes as shown in Table C.4.2. All volumetric laboratory tools were calibrated prior the cruise as stated in subsection (4.1). Then the actual concentration of nutrients in each fresh standard was calculated based on the ambient, solution temperature and determined factors of volumetric lab. wares. The calibration curves for each run were obtained using 4 levels, C-1, C-2, C-3 and C-4.

| Table C.4.1. Nominal concentrations of nutrients for A, B and C standards. Unit: μmol/l |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | C-1 | C-2 | C-3 | C-4 (Full scale) |
| NO3 | 26200 | 520 | LNSW* | 1/3 Full scale | 2/3 Full scale | 43.4 |
| NO2 | 12500 | 250 | LNSW* | 1/3 Full scale | 2/3 Full scale | 2.0 |
| PO4 | 2040 | 40.5 | LNSW* | 1/3 Full scale | 2/3 Full scale | 3.2 |
| Si | 35600 | 1950 | LNSW* | 1/3 Full scale | 2/3 Full scale | 155 |

| Table C.4.2. Working calibration standard recipes. |  |  |
| --- | --- | --- |
| C Std. | B-1 Std. | B-2 Std. |
| C-4 (Full scale) | 20 ml | 2 ml |
|  | LNSW* | C-4 (Full scale) |
| C-1 | 30 ml | 0 ml |
| C-2 | 20 ml | 10 ml |
| C-3 | 10 ml | 20 ml |

B-1 Std.: Mixture of nitrate, phosphate and silicate.

B-2 Std.: Nitrite.

LNSW*: 22 ml milli-Q water in 250 ml volumetric flask, and LNSW add to marked line.

(4.4) Renewal of in-house standard solutions

In-house standard solutions as stated in (4.3) were renewed as shown in Table C.4.3.

| Table C.4.3. Timing of renewal of in-house standards. |  |
| --- | --- |
| NO3, NO2, PO4, Si | Renewal |
| A-1 Std. (NO3) | no renewal |
| A-2 Std. (NO2) | no renewal |
| A-3 Std. (PO4) | no renewal |
| A-4 Std. (Si) | commercial prepared solution |
| B Std. |  |
| B-1 Std. | maximum 9 days |
| B-2 Std. | maximum 14 days |

| C Std. | Renewal |
| --- | --- |
| mixture of B-1 and B-2 Std. | Every measurement |

B-1 Std.: Mixture of nitrate, phosphate and silicate.

B-2 Std.: nitrite.

(5) Use of RMNS

The reference material of nutrients in seawater (hereafter RMNS), which was prepared by the General Environmental Technos Co. Ltd. (Kanso Technos), was used every analysis at each hydrographic station. According to Aoyama et al. (2010), the RMNS homogeneity is 0.1 % - 0.2 % in high concentration range, and stability is 48 - 71 months. By the use of RMNSs for the analysis of seawater, it is expected to secure stable comparability and uncertainty of data. If RMNS will be certified in the future, the traceability of our analysis value will be secured.

Aoyama et al. (2010) assigned nutrients concentrations for RMNS lot BA, AX, BE and AZ as shown in Table C.4.4.

| Table C.4.4. INSS assigned concentration of RMNSs. Unit: μmol/kg |  |  |  |
| --- | --- | --- | --- |
|  | Nitrate | Phosphate | Silicate |
| RMNS-BA | 0.07±0.01 | 0.061±0.007 | 1.61±0.07 |
| RMNS-AX | 21.44±0.05 | 1.614±0.006 | 58.05±0.12 |
| RMNS-BE | 36.70±0.04 | 2.662±0.005 | 99.20±0.08 |
| RMNS-AZ | 42.36±0.06 | 3.017±0.005 | 133.93±0.11 |

(5.1) RMNSs for this cruise

One hundred and five set of RMNS lots BA and BE were prepared to use every analysis at each hydrographic station. BA and BE were renewed every run. To check the inter-bottle consistency of RMNS, we re-measured the RMNS in the next analysis run. Sixteen of RMNS lots AX and AZ were prepared to use every 2 to 4 analysis and renewed every 2 or 3 runs in principle. The RMNS bottles were stored at a wet laboratory in the ship, where the temperature was maintained around 26 deg-C.

(5.2) Assigned concentration of RMNSs

We assigned nutrients concentrations for RMNS lots BA, AX, BE and AZ as shown in Table C.4.5 based on the analysis during the cruise. The measured concentration of RMNS lot BE during the cruise are shown in Figures C.4.6–C.4.8 as quality control charts. The concentration variations in these figures represent largely differences of the in-house standard. At Stn.25, one bottle of RMNS lot BE, No. 0138, showed obviously low concentrations. The measured values of all parameters were ca. 2% lower than expected. The concentrations of another bottle of RMNS lot BE, which was also analyzed in previous run, showed a good agreement with expected value. So, we neglected the result of No. 0138 of RMNS lot BE. At Stn.58, nitrate + nitrite concentration of BA No. 0756 had also shown unexpected low value, so it was neglected too.

The concentrations of RMNSs were in close agreement with expected values within the range of uncertainty except for the phosphate.

| Table C.4.5. Assigned concentration of RMNSs. Unit: μmol/kg |  |  |  |
| --- | --- | --- | --- |
|  | Nitrate + Nitrite | Phosphate | Silicate |
| RMNS-BA | 0.07±0.03 | 0.02±0.00 | 1.64±0.11 |
| RMNS-AX | 21.89±0.08 | 1.59±0.01 | 59.67±0.38 |
| RMNS-BE | 36.77±0.11 | 2.64±0.01 | 101.85±0.36 |
| RMNS-AZ | 42.42±0.13 | 2.98±0.01 | 137.29±0.29 |

Note: N(BA: Nitrate + Nitrite, Silicate)=104, N(BA: Phosphate)=105, N(BE: Nitrate + Nitrite, Phosphate)=104, N(BE: Silicate)=103, N(AX,AZ)=45.

Figure C.4.6. Result of RMNS lot BE concentrations of nitrate + nitrite during the cruise.

Figure C.4.7. Result of RMNS lot BE concentrations of phosphate during the cruise.

Figure C.4.8. Result of RMNS lot BE concentrations of silicate during the cruise.

(5.3) Relative standard deviation of RMNSs measurement

The relative standard deviation of lot BA, AX, BE and AZ throughout the cruise are shown in Table C.4.6.

| Table C.4.6. Relative standard deviation of RMNSs lot BA, AX, BE and AZ measurements in each run throughout cruise. |  |  |  |
| --- | --- | --- | --- |
|  | Nitrate + Nitrite CV % | Phosphate CV % | Silicate CV % |
| RMNS-BA | 41.03 | 16.02 | 6.47 |
| RMNS-AX | 0.35 | 0.43 | 0.31 |
| RMNS-BE | 0.30 | 0.30 | 0.19 |
| RMNS-AZ | 0.31 | 0.31 | 0.17 |

Note: N(BA: Nitrate + Nitrite, Silicate)=104, N(BA: Phosphate)=105, N(BE: Nitrate + Nitrite, Phosphate)=104, N(BE: Silicate)=103, N(AX,AZ)=45.

(6) Quality control

(6.1) Precision of nutrients analyses during the cruise

Precision of nutrients analyses during the cruise was evaluated based on 5 or 6 measurements of the C-4 (full scale) standard in each run. Summary of precisions are shown in Table C.4.7.

During this cruise, analytical precisions were 0.11 % for nitrate, 0.16 % for phosphate and 0.09 % for silicate in terms of mean of precision, respectively. The time series of precision are shown in Figure C.4.9 - C.4.11.

| Table C.4.7. Summary of precisions during the cruise. |  |  |  |
| --- | --- | --- | --- |
|  | Nitrate + Nitrite | Phosphate | Silicate |
|  | CV % | CV % | CV % |
| Median | 0.11 | 0.14 | 0.08 |
| Mean | 0.11 | 0.16 | 0.09 |
| Maximum | 0.32 | 0.38 | 0.31 |
| Minimum | 0.03 | 0.03 | 0.01 |
| Number | 105 | 105 | 104 |

Figure C.4.9. Time series of precision of nitrate + nitrite.

Figure C.4.10. Time series of precision of phosphate.

Figure C.4.11. Time series of precision of silicate.

(6.2) Replicate sample measurement

Replicate samples were analyzed at every hydrographic station. Total amount of the replicate sample pairs was 405. Summary of replicate sample measurements are shown in Table C.4.8, and Figure C.4.12 - C.4.14. During this cruise, the average difference and standard deviation of replicate measurement were 0.038±0.036 μmol/kg for nitrate + nitrite, 0.004±0.004 μmol/kg for phosphate and 0.140±0.139 μmol/kg for silicate, respectively.

| Table C.4.8. Average difference of replicate samples in each run throughout cruise. Unit: μmol/kg |  |  |
| --- | --- | --- |
| Nitrate + Nitrite | Phosphate | Silicate |
| 0.038±0.036 | 0.004±0.004 | 0.140±0.139 |

Note: N=403(nitrate, phosphate), N=399(silicate) at flag 2.

Figure C.4.12. Result of nitrate + nitrite replicate samplings (N=403) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of replicate samplings.

Figure C.4.13. Result of phosphate replicate samplings (N=403) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of replicate samplings.

Figure C.4.14. Result of silicate replicate samplings (N=399) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of replicate samplings.

(6.3) Duplicate sample measurement

Duplicate samples were analyzed at every hydrographic station. Total amount of the duplicate sample pairs was 295. Summary of duplicate sample measurements are shown in Table C.4.9, and Figure C.4.15 - C.4.17. During this cruise, the average difference and standard deviation of replicate measurement were 0.040±0.039 μmol/kg for nitrate + nitrite, 0.005±0.004 μmol/kg for phosphate and 0.167±0.163 μmol/kg for silicate, respectively.

| Table C.4.9. Average difference of duplicate samples in each run throughout cruise. Unit: μmol/kg |  |  |
| --- | --- | --- |
| Nitrate + Nitrite | Phosphate | Silicate |
| 0.040±0.039 | 0.005±0.004 | 0.167±0.163 |

Note: N=293(nitrate + nitrite), N=290(phosphate), N=289(silicate) at flag 2.

Figure C.4.15. Result of nitrate + nitrite duplicate samplings (N=293) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of duplicate samplings.

Figure C.4.16. Result of phosphate duplicate samplings (N=290) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of duplicate samplings.

Figure C.4.17. Result of silicate duplicate samplings (N=289) during RF10-05 against (a) station number, (b) sampling pressure, (c) concentration and (d) histogram of the result of duplicate samplings.

(7) Uncertainty

(7.1) Uncertainty associated with concentration level: Uc

The 44 sets of RMNS were analyzed during the cruise to make empirical equations to estimate uncertainty of concentrations of seawater samples throughout cruise. The average value and CV for each RMNS level were calculated, graphed, and a curve fit determined. The empirical equation (7.1) is an example of the curve fit between nutrients concentration Cx and the uncertainty at each concentration level.

Uncertainty for parameter X%=a+b1Cx+c1Cx2	-- (7.1)

Where Cx is concentration of sample for parameter X.

Empirical equations, eqs. (7.2), (7.3) and (7.4) were used to estimate uncertainty of measurement of nitrate + nitrite, phosphate and silicate during this cruise. The equations are based on analysis of 44 sets of RMNS lots BA, AX, BE and AZ. Figures C.4.18–C.4.20 show graphic presentations of eqs. (7.2) - (7.4).

Nitrate + Nitrite Concentration Cn in μmol/kg:

Uncertainty of measurement of nitrate (%) =

0.252 + 2.6214  (1/Cn) + 0.030  (1/Cn)2		-- (7.2)

Where Cn is nitrate concentration of sample.

Phosphate Concentration Cp in μmol/kg:

Uncertainty of measurement of phosphate (%) =

0.1588 + 0.4470  (1/Cp) + 0.0009  (1/Cp)2		-- (7.3)

Where Cp is phosphate concentration of sample.

Silicate Concentration Cs in μmol/kg:

Uncertainty of measurement of silicate (%) =

−0.092 + 43.838  (1/Cs) – 51.669  (1/Cs)2	-- (7.4)

Where Cs is silicate concentration of sample.

Figure C.4.18. Uncertainty of nitrate + nitrite concentration level.

Figure C.4.19. Uncertainty of phosphate concentration level.

Figure C.4.20. Uncertainty of silicate concentration level.

(7.2) Uncertainty of analysis between runs: Us

Uncertainty of analysis between runs (Us) was estimated from relative standard deviation of RMNS throughout cruise as shown in subsection (5.3).

(7.3) Uncertainty of analysis in a run: Ua

Uncertainty of analysis (Ua) was estimated from relative standard deviation of precision throughout cruise as shown in subsection (6.1).

(7.4) Conclusive uncertainty of nutrient measurements of samples: U

To determine the conclusive uncertainty of nutrient measurements of samples, we use two functions depending on Ua value acquired at each run as follows:

When Ua was small and measurement was well-controlled condition, the conclusive uncertainty of nutrient measurements of samples, U, might be as below:

U=Uc.							-- (7.5)

When Ua was relative large and the measurement might have some problems, the conclusive uncertainty of nutrient measurements of samples, U, can be expanded as below:

U=Uc2+Ua2. 					-- (7.6)

(8) Problems/improvements occurred and solutions

During the cruise, low-frequency noise (ca. 6 seconds per cycle) in the all channel output of AA III. So moving average was applied to all the raw data.

At Stn.16 (Lat. 30°39.21’N / Long. 136°59.69’E, RF3664), the silicate output of quality control samples and sensitivity compensation sample had exceeded the maximum value of the instrument setting. It was impossible to process silicate data for the station properly, so we neglect it. Pump tubes were replaced after the analysis.

Due to a problem on Phosphate data at Stn.36 (Lat. 20°59.83’N / Long. 136°58.21’E, RF3684), we had done another analytical run for the station. To reduce the analysis time, we omitted the C-2 and C-3 standard on these run, and processed the data as described below.

1. For each phosphate standard measurement in every run except for Stn.36, calculate the difference dCx between the concentration (Cx2) based on quadratic calibration equation and the concentration (Cx1) based on linear calibration equation.

2. Calculate mean and standard deviation of dCx at each level (C-1, C-2, C-3 and C-4) and reject the data for individual run if any of dCx at each level exceeds the range of mean±standard deviation.

3. The correction equation (8.1) is calculated by the regression analysis using all QCed pair of dCx and Cx1.

dCx(Cx1)=a+bCx1+cCx12	-- (8.1)

4. For phosphate of Stn.36, calculate tentative sample concentrations Ct1 based on linear calibration equation in each run.

5. The sample concentration C is obtained from equation (8.2) using Ct1 and eq. (8.1).

C=Ct1+dCx(Ct1)	-- (8.2)

The correction term dCx represents non-linearity of standard calibration. The correction for phosphate at Stn.36 was smaller than 0.003 μmol/kg.

(9) Results

(9.1) Comparison at cross-stations during this cruise

Cross-stations during this cruise were two stations. The one was located at 2°N/142°E, another was located 7°N/137°E. At stations of Stn.83 (RF3731) and Stn.104 (RF3755), hydrocast sampling for nutrients (nitrate, nitrite, phosphate, silicate) were conducted two times at interval of about five days. Each nutrients parameter profiles of the two hydrocasts agreed well within the range of uncertainty when correcting it by using RMNS. At stations of Stn.67 (RF3715), Stn.68 (RF3716) and Stn.124 (RF3772), hydrocast sampling for nutrients were conducted three times. Interval between the first and the second was about a week, interval between the second and the third was about two weeks. Each nutrients parameter profiles of the three hydrocasts agreed well within the range of uncertainty when correcting it by using RMNS. These profiles are shown in Figure C.4.21 - C.4.23.

Figure C.4.21. Comparison of nitrate +nitrite profiles between the first hydrocast (circle) and the second one (triangle) at the cross-stations of 2°N/142°E (left), and the first hydrocast (circle), the second one (triangle) and the third one (square) at the cross-stations of 7˚N/137˚E (right).

Figure C.4.22. Comparison of phosphate profiles between the first hydrocast (circle) and the second one (triangle) at the cross-stations of 2°N/142°E (left), and the first hydrocast (circle), the second one (triangle) and the third one (square) at the cross-stations of 7°N/137°E (right).

Figure C.4.23. Comparison of silicate profiles between the first hydrocast (circle)and the second one (triangle) at the cross-stations of 2°N/142°E (left), and the first hydrocast (circle), the second one (triangle) and the third one (square) at the cross-stations of 7°N/137°E (right).
(9.2) Comparison at cross-stations of WHP-P2 section in 2004 and WHP-P9 in 1994

We compared our nutrients data with gridded data of WHP-P2 at a cross point around 30°N/137°E. WHP-P2 line was observed two times, the repeat cruise was observed in 2004 by R/V Melville belonged to Scripps Institution of Oceanography (SIO). WHP-P9 line was observed in 1994 by JMA. These data may have inter-cruise differences because they did not measure the RMNS in their cruise. Summary of compared these data profiles shown in Figure C.4.24 - C.4.26.

Figure C.4.24. Comparison of nitrate + nitrite profiles at cross-station of WHP-P2. Circle, plus, square show the WHP-P9 in 1994 by JMA, WHP-P2 in 2004 by SIO and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.25. Comparison of phosphate profiles at cross-station of WHP-P2. Circle, plus, square show the WHP-P9 in 1994 by JMA, WHP-P2 in 2004 by SIO and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.26. Comparison of silicate profiles at cross-station of WHP-P2. Circle, plus, square show the WHP-P9 in 1994 by JMA, WHP-P2 in 2004 by SIO and WHP-P9 revisit in 2010 by JMA, respectively.

(9.3) Comparison at cross-stations of WHP-P3 section in 1985, 2005/06 and WHP-P9 in 1994

We compared our nutrients data with gridded data of WHP-P3 at a cross point around 24°N/137°E. WHP-P3 line was observed two times, the first was observed in 1985 by R/V Thomas G. Thompson belonged to SIO and the repeat cruise was observed in 2005/06 by R/V Mirai belonged to Japan Agency for Marine-Earth Science and Technology (JAMSTEC, 2007). WHP-P9 line was observed in 1994 by JMA. Our nutrients data at P9 revisit and JAMSTEC data in 2005/06 are comparable directly through the RMNS. However, SIO data in 1985 and JMA data in 1994 may have inter-cruise differences because they did not measure the RMNS in their cruise. Summary of compared these data profiles shown in Figure C.4.27 - C.4.29.

Note: Silicate data of WHP-P3 revisit (JAMSTEC, 2007) is corrected by a scale factor provided by M. Aoyama, PI of nutrients of the cruise (personal communication).

Figure C.4.27. Comparison of nitrate + nitrite profiles at cross-station of WHP-P3. Plus, square, triangle, circle show the WHP-P3 in 1985 by SIO, WHP-P9 in 1994 by JMA, WHP-P3 in 2005/06 by JAMSTEC and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.28. Comparison of phosphate profiles at cross-station of WHP-P3. Plus, square, triangle, circle show the WHP-P3 in 1985 by SIO, WHP-P9 in 1994 by JMA, WHP-P3 in 2005/06 by JAMSTEC and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.29. Comparison of silicate profiles at cross-station of WHP-P3. Plus, square, triangle, circle show the WHP-P3 in 1985 by SIO, WHP-P9 in 1994 by JMA, WHP-P3 in 2005/06 by JAMSTEC and WHP-P9 revisit in 2010 by JMA, respectively. Data of WHP-P3 revisit (JAMSTEC, 2007) is corrected by a scale factor provided by M. Aoyama, PI of nutrients of the cruise (personal communication).

(9.4) Comparison at cross-stations of WHP-P4 section in 1989, and WHP-P9 in 1994

We compared our nutrients data with gridded data of WHP-P4 at cross point around 9°N/137°E. WHP-P4 line was observed in 1989 by R/V Moan Wave belonged to University of Hawaii (UH). WHP-P9 line was observed 1994 by JMA. These data may have inter-cruise differences because they did not measure the RMNS in their cruise. Summary of compared these data profiles shown in Figure C.4.30 - C.4.32.

Figure C.4.30. Comparison of nitrate + nitrite profiles at cross-station of WHP-P4. Circle, plus, square show the WHP-P4 in 1989 by UH, WHP-P9 in 1994 by JMA and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.31. Comparison of phosphate profiles at cross-station of WHP-P4. Circle, plus, square show the WHP-P4 in 1989 by UH, WHP-P9 in 1994 by JMA and WHP-P9 revisit in 2010 by JMA, respectively.

Figure C.4.32. Comparison of silicate profiles at cross-station of WHP-P4. Circle, plus, square show the WHP-P4 in 1989 by UH, WHP-P9 in 1994 by JMA and WHP-P9 revisit in 2010 by JMA, respectively.

(10) Update of silicate data

May 27, 2018

For the silicate standard, we used “Silicon standard solution traceable to SRM from NIST SiO2 in NaOH 0.5 mol/l 1000 mg/L Si CertiPUR” provided by Merck with lot number HC814662. The silicate concentration is certified by NIST-SRM3150 as 999 mg/L with the expanded uncertainty of ± 5 mg/L (k=2). However, based on comparison with the solutions of lot HC074650 used in following cruise RF10-07, we assigned correction factor for the lot HC814662 as 0.975 to ensure internal comparability among our cruises in the future. This correction factor between the lots of HC814662 and HC074650 is consistent with that assigned by JAMSTEC (Uchida et al., 2015). Therefore, we updated silicate data to be multiplied by the factor 0.975.
References

Aoyama, M., S. Becker, K. Sato and D. Schuller (2009), Plan of use of RMNS during the CLIVAR P6 revisited cruise by R/V Melville. (unpublished manuscript).

Aoyama, M., A. G. Dickson, D. J. Hydes, A. Murata, J. R. Oh, P. Roose and E. Malcom. S. Woodward (2010), Comparability of nutrients in the world’s ocean, INSS international workshop 10-12 Feb. 2009, Paris

Armstrong, F. A. J., C. R. Stearns and J. D. H. Strickland (1967), The measurement of upwelling and subsequent biological processes by means of the Technicon TM Autoanalyzer TM and associated equipment, Deep-Sea Res., 14(3), 381–389.

Grasshoff, K., Ehrhardt, M., Kremling K. et al. (1983), Methods of seawater analysis. 2nd rev, Weinheim: Verlag Chemie, Germany, West.

JAMSTEC (2007), WHP P3 REVISIT DATA BOOK, WHP P03 REVISIT in 2005.

Murphy, J., and Riley, J.P. (1962), Analytica chimica Acta 27, 31-36.

Uchida, H., K. Katsumata, and T. Doi (2015): WHP P14S, S04I REVISIT IN 2012 DATA BOOK (http://www.godac.jamstec.go.jp/catalog/data/doc_catalog/media/
2012_WHP_P14S_S04I_all.pdf).

- Total Dissolved Inorganic Carbon (DIC) and Total Alkalinity (TA)

13 December 2013

(1) Personnel

Shinji MASUDA (GEMD/JMA)

Kazutaka ENYO (GEMD/JMA)

Naohiro KOSUGI (MRI/JMA)

(2) Overview

The concentration of total dissolved inorganic carbon (DIC) and total alkalinity (TA) were determined simultaneously from a single bottle of seawater sample by using two sets of custom-made DIC/TA analyzer manufactured by Nippon ANS Co. Ltd. (apparatus-A and apparatus-B). DIC was determined by coulometric analysis (Johnson et al., 1985, 1987) using an automated CO2 extraction unit and a coulometer (2009 model, Nippon ANS Co. Ltd.). TA was determined by one-step volumetric addition of hydrochloric acid (HCl) followed by the spectrophotometric analysis of pH with the sulfonephthalein indicator dye bromocresol green (Breland and Byrne, 1993) using an automated titration system equipped with CCD image sensor spectrophotometers (Hamamatsu, TA-CCD-A).

At each station, the precision of analysis was monitored using the Certified Reference Material (CRM) for DIC and TA (batches 93, 99 and 101) supplied by Dr. Andrew G. Dickson in Scripps Institution of Oceanography and a non-certified working reference material we have prepared from a surface seawater taken in the western North Pacific. A reference gas of 1.5% CO2 in air (Japan Fine Products) was also routinely measured to monitor the integrity of the coulometric cell assembly. However, it was found after the cruise that the elevated pressure in the head-space of coulometric cathode cell due to the insufficient venting capacity for carrier N2 stream had been making the DIC measurements less precise. It was also found that its effect was in part empirically corrected for by using the data of ambient pressure. We therefore standardized the DIC measurements with the analytical results of CRMs after pressure corrections and their certified DIC concentrations. Concentration of hydrochloric acid (HCl) in the titrant for TA analysis was also monitored by the analytical result of CRMs and their certified TA values. The overall precision of measurements as estimated from the replicate measurements of CRMs during the course of cruise were 3.2 mol kg-1 for DIC and 1.4 mol kg-1 for TA.

(3) Samplings

Measurements of DIC and TA in the full water column were made at a total of 53 stations (Leg 1: 31, Leg 2: 22) (Figure C.5.1). Intervals of sampling stations are 1° in latitude in open ocean zone and 15' to 40' in offshore regions near Japan and near Papua New Guinea.

Figure C.5.1. Station locations (left panel) and sampling layers of DIC and TA (middle and right panels)

Samples for the measurements of DIC and TA were drawn according to the procedures outlined by Dickson et al. (2007) from 10-L Niskin bottles into clean 300 cm3 Schott Duran® borosilicate glass bottles using silicone tubing. To minimize CO2 exchange with the ambient air, samples for DIC/TA were drawn next to those for CFCs and dissolved oxygen. Samples of near-surface seawater were collected from the underway seawater supply from the sea-chest (approx. 5 m).

Total of 176 pairs of replicate samples from the same Niskin bottle and total of 152 pairs of duplicate samples from different Niskin bottles tripped at the same depth were drawn for quality assurance of measurements.

Schott Duran® glass bottles were filled smoothly from the bottom after overflowing double a volume while taking care of not entraining any bubbles.

After creating 2 cm3 of headspace by removing sample to allow thermal expansion, the sample bottles were sealed with ground glass stoppers lubricated with Apiezon® grease (L). Until Stn.15, 0.2 cm3 of saturated mercury (II) chloride solution was added to the samples as a preservative. After Stn.18, no mercury (II) chloride was added and measurements of DIC and TA were started immediately after samplings. Samples were immersed in a thermo stated water bath (25.0 °C) for approx. 1 hour prior to analysis.

All the samples with the preservative were corrected to compensate dilution multiplied by 0.2/300 cm3/cm3 as a dilution factor to each concentration.

(4) Dissolved Inorganic Carbon (DIC)

(4-1) Instrumentation and procedures

The unit for DIC measurement in the coupled DIC/TA analyzer consists of a coulometer with a quartz coulometric titration cell (8 cm outer diameter), a CO2 extraction unit and a reference gas injection unit. The CO2 extraction unit includes a sample pipette (approx. 15 cm3) and a CO2 extraction chamber that is connected to a bottle of 10% v/v phosphoric acid and a carrier N2 gas supply, two thermoelectric cooling units and so on. The coulometric titration cell and the sample pipette are water-jacketed and are connected to a thermo stated (25 °C) water bath. The automated procedures of DIC analysis in seawater were as follows:

- Approximately 2 cm3 of 10% v/v phosphoric acid was injected to an “extraction chamber”, i.e., a glass tube (approx. 20 mm outer diameter. and 20 cm in length) with a course glass frit placed near the bottom. Purified N2 (Japan Fine Products, G1 grade >99.99995%) was then allowed to flow through the extraction chamber for 2 minutes to purge CO2 and other volatile acids dissolved in the phosphoric acid.

(b) A portion of sample seawater was delivered from the sample bottle into the sample pipette of CO2 extraction unit by pressurizing the headspace in the sample bottle. After temperature of the pipette was recorded, the sample seawater was transferred into the extraction chamber and mixed with phosphoric acid to convert all carbonate species to CO2 (aq).

(c) The acidified sample seawater was then stripped of CO2 with a stream of purified N2 (130 cm3 min-1) for 10 minutes. After being dehumidified in a series of two thermoelectric cooling unit (2 °C), the evolved CO2 in the N2 stream was introduced into the carbon cathode solution (UIC Inc.) in the coulometric titration cell where it reacts rapidly and quantitatively with 2-aminoethanol to form N-(2-hydroxyethyl)carbamic acid. This strong acid is immediately titrated with hydroxide ion (OH－) generated by the electrolysis of water on Pt electrode in the cathode.

Total amount of CO2 evolved from the acidified sample seawater is thus measured by integrating the total charge required to generate OH－ to neutralize the N-(2-hydroxyethyl)- carbamic acid. The end-point of titration is determined spectrophotometrically with the indicator dye thymolphthalein as the point of 30.0% transmittance of light at approx. 610 nm.

The entire sequence takes about 13 minutes for a sample. Once every 6 samples, additional 5 minutes are allowed for titration to evaluate the “background count level” of the coulometer.

Cathode and anode solutions of the coulometer were renewed at the beginning of DIC measurements at each station. After conditioning of the solutions, the amount of CO2 in 1.5% CO2 in air taken in an electro polished stainless-steel flask (approx. 60 cm3) was measured in order to monitor the integrity of the coulometric cell assembly. A bottle of CRM (batches 93, 99 and 101) was measured at each run of solutions. A working reference material (SSW-P) prepared from western North Pacific surface water was also measured at the beginning, in the middle, and at the end of measurements at a station.

(4-2) Calculation of DIC

Concentration of DIC (CT) in moles per kilogram of seawater (mol kg-1) was calculated from equation (1):

CT = NS / (c  VS  S)	,						(1)

where NS is the net counts of coulometer, c is the coulometer calibration factor, i.e., the counts of coulometer per mole of carbon, VS is the sample volume (volume of pipette), and S is the density of seawater that is calculated from the salinity of sample and its temperature in the pipette.

Net counts of coulometer, NS, is the counts for sample at 10 minutes after starting CO2 extraction/titration (N10) subtracted by background count level NB per 10 minute, i.e.,

NS= N10 – NB.							(2)

Background count levels were measured once every six sample measurements. We evaluated this using equation (3) from the increase of coulometer count from 10 minutes (N10) to 15 minutes (N15) after starting CO2 extraction/titration when CO2 in sample seawater is expected to have been completely evolved.

NB = (N15 – N10)  10 / (15－10)					(3)

In this cruise, NB values were averaged for measurements in each run of coulometric cathode and anode solutions.

A problem in DIC measurement in this cruise was that the coulometer counts for 1.5% CO2 in air, CRM and the working reference material were not as stable as we have expected before the cruise. For example, in the laboratory on land, coefficient of variation (C.V.) for NS of SSW measurements was usually less than 0.08% after NS was corrected for the changes in the density of SSW, but was 0.17% (411.46 ± 0.69 gC; N = 48) for apparatus A and 0.19% (411.78 ± 0.78 gC; N = 66) for apparatus B. After the cruise, it was found that the headspace in the cathode cell have been pressurized by the stream of N2 due to its insufficient venting capacity, and this had been making the coulometer count less stable. It was also found that the change in the coulometer count due to the headspace-pressurizing can be in part empirically corrected for by the change in the ambient pressure (P) as:

NS’ = NS / {1 – 0.5  (1 – P/1013.25hPa)}				(4)

We used the value of NS’, instead of NS, for the calculation of DIC concentration using equation (1). The decreasing trend of NS was decreased by using the value NS’ (Figure C.5.2).

Figure C.5.2. Coulometer counts NS for SSW (left panel) and the counts corrected for the changes in ambient pressure, NS’ (right panel). The colour of plots denotes the apparatus-A (red) and apparatus-B (blue), respectively.

(4-3) Results of DIC measurements of CRM

The value of c Vs was determined for each DIC measurement of CRM (Batch 93, 99 and 101) from equation (5):

c Vs = NS’ / (CT, CRM  25), 					(5)

where CT, CRM denotes the certified DIC concentration of CRM. The results are shown in Figure C.5.3. Values of cVs were averaged over the same leg for each apparatus, except at station 1 where measures were made using the apparatus-B, and used for the calculations of DIC in sample seawaters (Table C.5.1).

Standard deviation of the differences in the analytical DIC concentration of CRM calculated with the averaged c Vs was ±1.5 mol kg-1 (N = 25) (Figure C.5.4a), and the repeatability as estimated from the absolute difference in the replicate analyses of CRM was ±2.8 mol kg-1 (Figure C.5.4b). The precision thus evaluated from the measurement of CRM was ±3.2 (= (1.52 + 2.82)1/2) mol kg-1.

Standard deviation of the DIC concentration of working reference material (SSW-P) was ±3.4 mol kg-1 (N = 113) (Figure C.5.5). It is consistent with repeatability of DIC measurements as evaluated from the measurements of the CRMs.

Figure C.5.3. Calibration factor c  Vs of the coulometeric systems determined from the measurements of CRM batch 93 (diamond), batch 99 (triangle) and batch 101 (inverted triangle) for apparatus A (red) and apparatus B (blue). Horizontal lines show the averages of c･Vs used for the calculation of DIC concentration in seawater samples (Table C.5.1).

Table C.5.1.  Summary of calibration factor, cVs.

| Apparatus | Leg | Mean ± sd. |
| --- | --- | --- |
| A | 1 | 0.19909 ± 0.00010 |
| A | 2 | 0.19924 ± 0.00021 |
| B | 1*1 | 0.20047 ± 0.00006 |
| B | 1*2 | 0.19887 ± 0.00022 |
| B | 2 | 0.19894 ± 0.00005 |

*1 For station Stn.1. *2 For station Stn.9 and later.

| (a) |  |
| --- | --- |
| (b) |  |

Figure C.5.4. (a) Differences in the measured DIC concentration from the certified value and (b) the absolute differences in replicate DIC measurements (R-chart) of CRM batches 93 (diamond), 99 (triangle) and 101 (inverted triangle) for apparatus-A (red) and apparatus-B (blue). Control limits in (a) were defined from the average (m) and the standard deviation (sd) of the differences; UCL = m + 3 sd, UWL = m + 2 sd, LCL = m – 3 sd, and LWL = m – 2 sd. UCL and UWL in (b) were defined from mean difference (R) as 3.167 R and 2.512 R, respectively (Dickson et al., 2007).

Figure C.5.5 Results of the measurements of working reference materials (SSW-P).

(5) Total Alkalinity

(5-1) Instrumentation and procedures

TA is measured by one-step volumetric addition of hydrochloric acid (HCl) to a known amount of sample seawater with prompt spectrophotometric measurement of excess acid using the sulfonephthalein indicator bromocresol green (BCG) (Breland and Byrne, 1993). The unit for TA measurements in the coupled DIC/TA analyzer consists of sample treatment unit (Nippon ANS) with a calibrated sample pipette and an open titration cell that are water-jacketed and connected to a thermo stated (25 °C) water bath, an auto syringe (PSD/8, Hamilton) connected to a bottle (Schott Duran®, 1 dm3) of titrant stored at 25 °C, and a double-beam spectrophotometric system with two CCD image sensor spectrometers (C10083CAH, Hamamatsu Photonics) combined with a high power Xenon lamp (HPX-2000, Ocean Optics). The mixture of 0.05N HCl and 40 mol dm-3 BCG in 0.65M NaCl solution was used as titrant to automatically titrate the sample as follows:

(a) A portion of sample seawater was delivered into the sample pipette (approx. 42 cm3) after the other portion is delivered into the DIC unit for a DIC measurement. After the temperature of pipette was recorded, the sample was transferred into a cylindrical quartz cell (4 cm o. d.).

(b) An absorption spectrum of sample seawater in the visible light domain was then measured, and the absorbances at wavelengths of 444 nm, 509 nm, 616 nm and 730 nm as well as the temperature in the cell were recorded.

(c) The titrant that includes 0.05N HCl was added to the sample seawater by using the auto syringe so that pH of sample seawater becomes in the range between 3.85 and 4.05 after the next step (d).

(d) While the acidified sample was being stirred, the evolved CO2 was purged with the stream of purified N2 bubbled into the sample at approx. 100 cm3 min-1 for 5 minutes.

(e) After leaving the bubbled acidified sample still for 1 minute, the absorbance of bromocresol green in the sample was measured in the same way as described in (b), and pH of the acidified seawater was precisely determined spectrophotometrically.

A typical titration including rinse, fill and discharge takes about 13 minutes.

The data of absorbance (A) and temperature (T) were processed to calculate the concentration of excess acid:

pHT = – log10([H＋]T/mol kg-seawater-1)

= 4.2699 + 0.02578(35 – S) + log{(R25 – 0.00131)/(2.3148 – 0.1299R25)}

– log(1 – 0.001005  S) 						(6)

R25 = RT  {1 + 0.00907  (25 – (T/ ºC))}				(7)

RT=A616SA-A616S-A730SA+A730SA444SA-A444S-A730SA+A730S 	(8)

S is salinity of sample in PSS-78 that was measured separately. AλS and AλSA denotes absorbance of seawater and acidified seawater, respectively, at wavelength  nm.

The concentration of excess acid [H＋]T determined is then combined with the volume of sample seawater (VS / dm-3), the volume of titrant (VA / dm-3) added to the sample, and molarity of hydrochloric acid (MA / mol dm-3) in the titrant to calculate total alkalinity (AT) in the unit of mol kg-sewater-1:

AT = (–[H+]T  (VS + VA) SA + MA  VA) / (VS  S)			(9)

S and SA denotes the density of seawater sample before and after the addition of titrant, respectively. Here we assumed that SA is equal to S, since the density of titrant has been adjusted to that of seawater by adding sodium chloride and the volume of titrant (approx. 2.5 cm3) is no more than approx. 6% of seawater sample.

(5-2) Volume of sample seawater, VS

The volumes of sample seawater, VS, i.e., the volume of pipette in the TA measurement unit of DIC/TA analyzer, was calibrated gravimetrically and summarized in Table C. 5.2.

Table C.5.2 Summary of sample volumes of seawater VS for TA measurements.

| Apparatus | Leg | Volume (ml) |
| --- | --- | --- |
| A | 1-2 | 42.050 |
| B | 1 | 43.354 |
| B | 2 | 41.967 |

(5-3) Preparation of titrant (0.05N HCl)

76 g of sodium chloride (NaCl) (Wako Pure Chemical Industries, Ltd.) and 0.12 g of bromocresol green (BCG) (Acros Organics) were dissolved in 200 cm3 of 0.5 mol dm-3 HCl solution (Wako Pure Chemical Industries, Ltd.). The solution was diluted with deionised water to a final volume of 2 dm3 at 25 °C. The concentration of HCl, NaCl and BCG was 0.05 mol dm-3, 0.65 mol dm-3 and 40 mol dm-3, respectively. Sodium chloride was added to make the density and the ionic strength of the solution close to those of seawater.

(5-4) Results of TA measurements of CRMs.

Measurements of TA for CRMs (batch 93, 99 and 101) were made and the apparent molarity of hydrochloric acid (MA / mol dm-3) in the titrant was determined from equation (10):

MA = (AT, CRM  VS  S + [H+]T  VS  SA ) / VA			(10)

The analytical results of MA was averaged for each bottle of titrant unless the drift of molarity was clearly seen (Figure C.5.6), and was used to calculate the TA in sample seawaters.

Standard deviation of the differences in the analytical results of TA for CRMs calculated with the averaged MA was ±0.8 mol kg-1 (N=27) (Figure C.5.7a). The repeatability as estimated from the absolute difference in the replicate analyses of CRMs was ±1.2 mol kg-1 (Figure C.5.7b). The precision thus evaluated from the measurement of CRM was ±1.4 mol kg-1.  Standard deviation of TA of working reference material (SSW-P) was ±3.2 mol kg-1 (N=114) (Figure C.5.8).

Figure C.5.6. Concentration of acid in HCl solution. The notations of the plots were same as those in Figure C.5.4. Line indicates the mean of each bottle of HCl solution. The lot of HCl solution was switched at the stations indicated as vertical dotted lines.

| (a) |  |
| --- | --- |
| (b) |  |

Figure C.5.7. (a) Differences in the measured TA from the certified values and (b) the absolute differences in replicate DIC measurements (R-chart) of CRM batches 93 (diamond), 99 (triangle) and 101 (inverted triangle) for apparatus-A (red) and apparatus-B (blue).  Control limits and warning limits were defined in the same way as in Figure C.5.4.

Figure C.5.8. Results of TA of working reference material (SSW-P).

(6) Assignment of quality flag

Quality code was assigned for each of DIC and TA measurements according to the WHP quality code definitions for water sample measurements (Swift and Diggs, 2008, Swift, 2010). Summary of assigned quality flags is shown in Table C.5.3. Data from replicate samples were averaged and flagged 6 if both flags have been assigned 2. If either of flags has been assigned 3 or 4, younger flag was selected.

TA data of samples collected at stations and layers listed in Table C.5.4 were assigned questionable (QF=3). A possible reason for these questionable data was insufficient rinsing of the optical titration cell after a previous measurement due to the malfunction of peristaltic pump to supply rinsing water to the cell.

Table C.5.3. Summary of assigned quality flags.

| Flag | Definition | DIC | TA |
| --- | --- | --- | --- |
| 2 | Good | 1874 | 1824 |
| 3 | Questionable | 118 | 155 |
| 4 | Bad (Faulty) | 20 | 36 |
| 5 | Not reported | 3 | 0 |
| 6 | Replicate measurements | 144* | 152* |
| Total number of samples |  | 2015* | 2015* |

*Samples of flag 6 are counted as flag 2.

Table C.5.4. Questionable measurements of TA.

| Station | Depth |
| --- | --- |
| Stn.9 Stn.11 Stn.26 Stn.28 Stn.40 | Deeper than or equal to 500 m Shallower than or equal to 1130 m Deeper than or equal to 1330 m Shallower than or equal to 100 m Shallower than or equal to 730 m |

(7) Results of replicate and duplicate sample measurements

Total of 176 pairs of replicate samples, drawn from a same Niskin bottle, were collected at every stations for DIC and TA measurements. The average of difference in acceptable pairs of measurements was 1.8 μmol kg-1 for DIC (N=144) and 1.9 μmol kg-1 for TA (N=152).  Total of 152 pairs of duplicate samples, drawn from a different Niskin bottle tripped at the same depth, were also collected at every stations for DIC and TA measurements. The average of difference in acceptable pairs of measurements was 2.0 μmol kg-1 for DIC (N=137) and 1.7 μmol kg-1 for TA (N=133).

Summary of replicates are shown in Figures C.5.9 and C.5.10 and in Table C.5.5. Summary of duplicates are shown in Table C.5.6. The average and standard deviation were calculated using a procedure described in SOP23 in Dickson et al. (2007).

Figure C.5.9. Absolute differences in DIC between measurements of replicate samples (N=144) during the cruise RF10-05 against (a) station number, (b) sampling depth, and (c) concentration of DIC. The histogram of absolute difference is shown in (d).

Figure C.5.10. Same as Figure C.5.9 but for TA.

Table C.5.5 Summary for the measurements of replicate samples.

|  | DIC (mol/kg) | TA (mol/kg) |
| --- | --- | --- |
| Total number of good measurement pairs | 144 | 152 |
| Average of absolute difference | 1.8 | 1.9 |
| Standard deviation of measurements | 1.7 | 1.7 |

Table C.5.6. Summary for the measurements of duplicate samples.

|  | DIC (mol/kg) | TA (mol/kg) |
| --- | --- | --- |
| Total number of good measurement pairs | 137 | 133 |
| Average of absolute difference | 2.0 | 1.7 |
| Standard deviation of duplicate sampling | 1.7 | 1.6 |

(8) Comparisons at cross-over stations within this cruise

There were two cross-over stations that were occupied multiple times within this cruise. The one was located at 7˚N, 137˚E. This station was occupied three times; station 67 (RF3715) on July 23, station 68 (RF3716) on August 2, and station 124 (RF3772) on August 16. The other was located at 2˚N, 142˚E. This station was occupied twice; station 83 (RF3731) on August 6 and station 104 (RF3755) on August 11. Vertical profiles of DIC and TA at these stations are shown in Figures C.5.11 and C.5.12.

Figure C.5.11. Vertical profiles of DIC at the cross-over stations at 2˚N, 142˚E (left panel: station 83 on August 6 (cross) and station 104 on August 11 (circle)) and at 7˚N, 137˚E (right panel: station 67 (cross) on July 23, station 68 (triangle) on August 2, and station 124 (RF3772) on August 16 (circle)).

Figure C.5.12. Same as Figure C.5.11 but for TA.

(9) Comparisons at cross-over stations with other WHP and its revisit cruises

The WHP section P9 has been observed in 1994 (Kaneko et al., 1998). Section P9 intersects section P2 at 30˚N, 137˚E and section P3 at 24˚N, 137˚E. DIC and TA in section P2 have been observed twice in the past; first by the cruise 49K6KY9401_1 of R/V Kaiyo-Maru in 1994 and second by the cruise 318M200406 of R/V Melville in 2004, and those in section P3 has been observed by the cruise 49NZ20051127 of R/V Mirai in 2005/2006. Summary of the comparisons of vertical profiles at cross-over stations are shown in Figure C.5.13 and C.5.14.

Figure C.5.13. Vertical profiles of DIC with WHP P2 (left panel: P9 at 1994 (square), P2 at 1994 (triangle), P2 at 2004 (cross) and P9 at 2010 (circle)) and WHP P3 (right panel: P3 at 2005/06 (square) and P9 at 2010 (circle)) at cross-over stations.

Figure C.5.14. Same as Figure C.5.13 but for TA.

References

Breland II, J. A. and R. H. Byrne (1993): Spectrophotometric procedures for determination of sea water alkalinity using bromocresol green, Deep-Sea Res. I, 470, 629-641.

Dickson, A. G., Sabine, C. L. and Christian, J. R. (Eds.) (2007): Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.

Johnson, K. M., A. E. King and J. McN. Sieburth (1985): Coulometric TCO2 analyses for marine studies; an introduction. Marine Chemistry, 16, 61-82.

Johnson, K. M., J. McN. Sieburth, P. J. IeB Williams and L Brändström (1987): Coulometric total carbon dioxide analyses for marine studies: Automation and calibration. Marine Chemistry, 21, 117-133.

Kaneko, I., Y. Takatsuki, H. Kamiya and S. Kawae (1998): Water property and current distributions along the WHP-P9 section (137°-142°E) in the western North Pacific. J. Geophys. Res., 103, 12959-12984.

Swift, J. H. and S. C. Diggs (2008): A Guide to Submitting CTD/Hydrographic/Tracer Data and Associated Documentation to the CLIVAR and Carbon Hydrographic Data Office (Appendix C: Description of WHP-Exchange Format for CTD/Hydrographic Data)., http://cchdo.ucsd.edu/policies.html.

Swift, J. H. (2010): Reference-quality water sample data: Notes on acquisition, record keeping, and evaluation. IOCCP Report No.14, ICPO Pub. 134, 2010 ver.1

- pH

13 December 2013

(1)  Personnel

Shinji MASUDA (JMA)

Kazutaka ENYO (JMA)

Naohiro KOSUGI (MRI)

(2)  Station occupied

A total of 53 stations (Leg 1: 31, Leg 2: 22) were occupied for hydrogen ion index (pH). Station location and sampling layers of pH are shown in Figure C.6.1.

Figure C.6.1. Station location (left panel) and sampling layers of pH (right panels)

(3)  Method

(3.1)  Principle

The pH analysis was made using spectrometry of indicator dye m-cresol purple (Saito et al., 2008; Clayton and Byrne, 1993). The pH was reported as the value at temperature of 25°C in “total hydrogen ion scale”. In order to state clearly the scale of pH, we mention hereafter as “pHT” that is defined by equation (1),

pHT=-log10[H+]TC0	---(1)

where, [H+]T denotes the concentration of hydrogen ion expressed in the total hydrogen ion scale: [H+]T=[H+]F1+[SO4]TKHSO4-, where, [H+]F is the concentration of free hydrogen ion, [SO4]T is the total concentration of sulphate ion and KHSO4- is acid dissociation constant of hydrogen sulphate ion (Dickson, 1990). C0 in equation (1) is the standard value of concentration (1 mole per kilogram of seawater, mol kg–1).

(3.2)  pHT Reagents

- m-Cresol purple solution

The air in a borosilicate glass flask (2 dm3) was replaced by pure nitrogen. 0.67 g of m-cresol purple sodium salt (pure water soluble, 199250050, ACROS) was dissolved in 1 kg of deionised water prepared with water purifier “Autopure WR700” (Yamato Scientific Co. Ltd.). Small amount of diluted NaOH (approx. 0.25 mol dm–3) solution was added to regulate the pH (free hydrogen ion scale) of indicator solution to 7.9 ± 0.1. The pH of indicator solution was monitored using glass electrode pH meter.

(3.3)  Instruments and procedure

Custom-made pHT analysers (2009 model; Japan ANS Co., Ltd, Tokyo, Japan) was prepared and operated in the cruise. The analyser comprised of a sample dispensing unit (box 1 in Figure C.6.2), a pre-treatment unit combined with an automated syringe (PSD/8, HAMILTON; box 2 in Figure C.6.2), and two (sample and reference) spectrophotometers (C10083CAH, Hamamatsu Photonics) combined with a high power xenon light source (HPX-2000, Ocean Optics) (box 3 in Figure C.6.2). Whole analyser system was controlled by a personal computer (EX/522PDET3, TOSHIBA) using a custom-made software (Japan ANS Co., Ltd) that runs on an operating system Windows XPTM.

The sample dispensing unit had an auto-sampler (6 ports), which took a seawater from 250 cm3 borosilicate glass bottles (017030-250, Shibata) to the pre-treatment unit. Spectrophotometric cell (4 in Figure C.6.2) was made of quartz tube (inner diameter: 4 mm; outer diameter: 6 mm; length: approx. 30 cm) that has figure of “U”. This cell was covered with stainless bellows tube to keep the external surface dry and for total light to reflect in the tube. The temperature of the cell was regulated to 25.0 ± 0.1°C by means of immersing the cell into the thermostat bath, where the both ends of bellows tube located above the water surface of the bath. Spectrophotometer, cell and light source were connected with optical fibre (red lines in Figure C.6.2).

3

4

1

2

Figure C.6.2. Diagram of pHT measurement apparatus. 1: Sample dispensing unit, 2: automated syringe to inject indicator, 3: two (sample and reference) spectrophotometers with high power Xe light source and 4: spectrophotometric cells.

The analysis procedure was as follows:

a) Seawater was ejected from a sample loop (thick line in Figure C.6.2)

b) A portion of sample (approx. 30 cm3) was introduced into a sample loop including spectrophotometric cell. The spectrophotometric cell was flushed two times with sample in order to remove air bubbles.

c) An absorption spectrum of seawater in the visible light range was measured. Absorbance at wavelengths of 434 nm, 488 nm, 578 nm and 730 nm as well as cell temperature were recorded. To eject air bubbles from the cell, the sample was moved four times (approx. 30 cm each in loop tube whose inner diameter was 1/16 inch) and the absorbance was recorded at each stop.

d) 80 l of indicator m-cresol purple solution was injected to the loop.

e) Circulating 2 minutes 40 seconds through the loop tube, seawater sample and indicator dye was mixed together. The final m-cresol purple concentration in the sample was approx. 4 mol dm–3.

f) Absorbance of m-cresol purple plus seawater was measured in the same way described above (c).

(4)  Seawater Sampling

Samples for pHT analysis were drawn with the similar way of dissolved inorganic carbon (see chapter C05) from 10-liter Niskin bottles into clean 250 cm3 borosilicate bottles (Shibata) using silicone rubber tubing on the petcock. To avoid contamination from the air, samples for pHT were drawn next to the sampling of dissolved inorganic carbon (DIC) and total alkalinity (ALK). Surface sample was collected from continuous flow line pumped from sea-chest. In order to avoid CO2 exchange with the air, the end of tubing was inserted to the bottom of the bottle, and then the sample was dispensed smoothly into a bottle. Sample of approximately double the volume of bottle was overflowed. The bottle was plugged temporally with a ground glass stopper.

After sampling, 2 cm3 of seawater was removed from bottle to allow thermal expansion of sample. Until the station RF3663, 0.2 cm3 of saturated mercury (II) chloride solution was added to prevent change in pHT caused by biological activity. After the station RF3666, pHT was analysed immediately after sampling instead of adding mercury chloride solution to the sample seawater.

Bottle was sealed with greased (Apiezon-L) ground glass stopper. Sample bottles were stored at room temperature while awaiting analysis. Sample bottles had been immersed in isothermal bath to keep in 25.0 °C for over 1 hour before analysis.

(5)  pHT measurement

(5.1)  Calculation of pHT from measured absorbance

pHT was calculated from the measured absorbances based on the following equations (2) – (4).

pHT=pK2+log10[I2-][HI-]

=pK2+log10R-0.006912.2220-0.1331R -- (2)

R=A578SD-A578S-A730SD+A730SA434SD-A434S-A730SD+A730S  -- (3)

where pK2 is the acid dissociation constant of m-cresol purple, [I2–] / [HI–] is the ratio of m-cresol purple base form (I2–) concentration over acid form (HI–) concentration, which is estimated from absorbance ratio R and the ratios of extinction coefficients (Clayton and Byrne, 1993). AλS and AλSD in equation (3) are absorbances of seawater itself and dye plus seawater, respectively, at wavelength  (nm). The value of pK2 (=-log10K2k0, k0 = 1 mol / kg) had also been expressed as a function of temperature T (in Kelvin) and salinity S (in psu) by Clayton and Byrne (1993), but the calculated value has been subsequently corrected by 0.0047 on the basis of a reported pHT value accounting for “tris” buffer (DelValls and Dickson, 1998):

pK2=pK2Clayton & Byrne, 1993+0.0047

=1245.69T+3.8322+0.0021135-S.	--(4)

293 K ≤ T ≤ 303 K, 30 ≤ S ≤ 37

(5.2)  pHT Perturbation caused by addition of dye solution

The injection of m-cresol purple solution affects the pHT of seawater sample because the acid base equilibrium of the seawater is disrupted by the addition of the dye acid-base pair (Dickson et al., 2007). We corrected the R in Equation (2) for the perturbation using empirical method (Equation (5)) in which a second aliquot of dye solution is added to the seawater sample (Dickson et al., 2007; Clayton and Byrne, 1993).

R = R1 – R,

R = R2 – R1 = R1 – R (Assumption),	--(5)

where, R1 and R2 are the absorbance ratio after the addition of first and second aliquot of dye solution, respectively. The value of R depended on the pHT of sample. We expressed R as a quadratic function of R1 based on experimental R2 – R1 data obtained at this cruise (Figure C.6.3). 6 samples at each station were analysed to obtain R2 – R1 data, R was expressed as a quadratic function of R1 and the pHT was evaluated from R = R1 – R using equation (2).

Figure C.6.3. pHT perturbation caused by the addition of indicator dye solution. The perturbation was expressed as a difference in absorbance ratio between first and second aliquot, R = R2 – R1. The colour of plots denotes the station numbers; red: RF3649 - RF3714, black: RF3716 – RF3731, red: RF3742 – 3772. The result of quadratic regression was listed in Table C.6.1.

Table C.6.1. The coefficients of quadratic regression of the pHT perturbation that was caused by the addition of m-cresol purple solution and was expressed as the difference in absorbance ratio R. ∆R=C2×R12+C1×R1+C0

| Stations | C2 | C1 | C0 |
| --- | --- | --- | --- |
| RF3649 - RF3714 RF3716 – RF3731 RF3742 – RF3772 | –3.53 × 10–3 –7.73 × 10–3 –5.27 × 10–3 | – 1.19 × 10–2 – 3.16 × 10–4 – 6.98 × 10–3 | 0.0156 0.0091 0.0125 |

(6) Quality assurance

(6.1) CRM and in-house standard seawater measurements

To check the repeatability and/or reproducibility of measurements, we analysed two batches (98 and 100) of certified reference materials (CRMs) that were prepared by Dr. A.G. Dickson at Scripps Institute of Oceanography. Three samples were extracted and analysed from a bottle and the difference in initial two data were plotted (R-chart, Figure C.6.4). The upper control limit (UCL), the upper warning limit (UWL) and the standard deviation were estimated using equations described in SOP22 of Dickson et al. (2007). The standard deviation estimated from accepted data (less than UCL) was 0.0010 for batch 98 and 0.0013 for batch 100, respectively.

Figure C.6.4. Absolute differences in replicate measurements of pHT in CRM for each bottle (R-chart). The colour of the plots denotes the batch number; blue: batch 98, red: batch 100.

| a) |  |
| --- | --- |
| b) |  |

Figure C.6.5. a) Measured pHT of CRMs, b) Differences in pHT of CRMs between measured and reference values that was calculated from certified values of DIC and ALK using acid dissociation constants of carbonates described by Lueker et al. (2000). The colour of plots denotes the batch of CRM; blue: batch 98, red: batch 100. UCL and LCL are upper and lower control limit, respectively, that is defined as mean ± 3 . UWL and LWL are upper and lower warning limit, respectively, that are defined as mean ± 2 .

All results of CRM measurements were plotted in Figure C.6.5 a). Measured pHT was compared with “reference value” that was calculated from certified values of DIC and ALK using acid dissociation constants of carbonic acid and hydrogen carbonate ion described by Lueker et al. (2000). The concentrations of phosphate and silicate were also used for calculation, although the values are not certified. The difference in measured and reference values were plotted in Figure C.6.5 b). The offset of measured value was –0.0010 ± 0.0024 (mean ± standard deviation, n = 85).

In order to monitor the condition of apparatus, in-house standard seawater (SSW) batch J, which was made using the CRM’s manner, was measured. The mean and standard deviation was 7.9392 ± 0.0028 (n = 97) (Figure C.6.6).

Figure C.6.6. Measured pHT of in-house SSW.

(6.2) Repeatability of water column samples measurements

To check the repeatability of water column samples measurements, we measured replicate and duplicate samples. At each hydrographic station where sea water samples were drawn, two or more Niskin bottles were closed at the same layer (“Duplicate” sampling) if Niskin bottles on carousel sampler are more than the layers to be sampled at the station. A couple of samples were drawn from each Niskin bottle of 3 layers to obtain 3 “Replicate” samples at each hydrographic station.

Total amounts of the duplicate and replicate sample pairs were 162 and 175, respectively. The control limits and standard deviations were calculated using the method described in SOP22 of Dickson et al. (2007). The average differences and the standard deviations estimated from accepted data are listed in Table C.6.2.

Table C.6.2 Summary of replicate and duplicate samples measurements

|  | Duplicate | Replicate |
| --- | --- | --- |
| Total number of good measurement pairs | 152 | 163 |
| Average of absolute difference | 0.0019 | 0.0021 |
| Standard deviation of measurements | 0.0017 | 0.0018 |

(6.3)  Quality flag assignment

Summary of assigned quality control flags are listed in Table C.6.3. The replicate data were averaged and flagged 6 if both of the flag were 2. If either of the flag was 3 or 4, younger flag was selected.

Table C.6.3 Summary of assigned quality control flags.

| Flag | Definition | Number of samples |
| --- | --- | --- |
| 1 | need further correction | 1840 |
| 3 | Questionable | 98 |
| 4 | Bad (Faulty) | 21 |
| 5 | Not reported | 3 |
| 6 | Replicate measurements | 163* |
| Total number of samples |  | 1962* |

* Samples of flag 6 are counted as flag 1

Figure C.6.7. Result of pHT replicate samplings (n=163) during the cruise RF10-05 versus (a) Station number, (b) Sampling depth, (c) pHT values and (d) Histogram of the result of replicate samplings. The lines in the panel (a) indicate upper control limit (thick), upper warning limit (dashed) and average of absolute difference (dotted), respectively.

Figure C.6.8. Result of pHT duplicate samplings (N=152) during the cruise RF10-05 versus (a) Station number, (b) Sampling depth, (c) pHT values and (d) Histogram of the result of duplicate samplings.

(7)  Problems

pHT data of 1840 samples (Table C.6.3) were assigned as QF.1 because the data could have significant bias due to non- linear response of our spectrophotometer. A method of correction is under consideration.

pHT data of 98 samples (Table C.6.3) were assigned as questionable data (QF = 3) because the data was sufficiently (approx. 2  of measurements) apart from smooth line of vertical profile of a station. The pHT data for these samples showed different characters from the data of DIC, ALK, dissolved oxygen or nutrients. Possible reason for this problem was a lack of rinsing optical cell after a measurement, because of malfunction of circulation pump of pre-treatment unit. pHT data of 21 samples (Table C.6.3) were assigned as bad (faulty) data (QF = 4). The absorbance at wavelength 730 nm of these data was anomalously high, which represented air bubbles remained in an optical cell. If all of five absorbance data per a sample measurement was faulty, the data was assigned as bad. pHT data of 3 samples were not reported (QF = 5) because the samples were lost in the measurement procedure, for example, attaching sample bottles at the wrong position of sampler.

(8)  Results

(8.1)  Comparison at cross-stations during the cruise

There were two cross-stations during the cruise. The one was located at 2˚N/142˚E, another was located at 7˚N/137˚E. At stations of Stn.83 (RF3731) and Stn.107 (RF3755), hydrocast sampling for pHT was conducted two times at interval of about five days.

At stations of Stn.67 (RF3715), Stn.68 (RF3716) and Stn.124 (RF3772), hydrocast sampling for pHT was conducted three times. Interval between the first and the second was about a week, interval between the second and the third was about two weeks.

These profiles are shown in Figure C.6.9.

Figure C.6.9. Comparison of pHT profiles between the first hydrocast (cross) and the second one (circle) at the cross-stations of 2˚N/142˚E (left panel), and between the first hydrocast (cross), the second one (triangle) and the third one (circle) at the cross-stations of 7˚N/137˚E (right panel)

(8.2)  Comparison at cross-stations of WHP-P3 section

We compared our pHT data and WHP-P3 at a cross point (around 24˚N/137˚E). WHP-P3 line was observed in 2005/06 by R/V Mirai that belonged to Japan Agency for Marine-Earth Science and Technology (JAMSTEC).

Summary of the comparison of these profiles is shown in Figure C.6.10.

Figure C.6.10. Comparison of pHT profiles at cross-stations of WHP-P3. Circle show this cruise.

References

Clayton T.D. and R.H. Byrne 1993. Spectrophotometric seawater pH measurements: total hydrogen ion concentration scale calibration of m-cresol purple and at-sea results. Deep-Sea Res. I, 40, 2115-2129.

Dickson, A.G. 1990. Standard potential of the reaction: AgCl(s) + 1/2 H2(g) = Ag(s) + HCl(aq), and the standard acidity constant of the ion HSO4– in synthetic sea water from 273.15 to 318.15 K. J. Chem. Thermodynamics, 22, 113-127.

Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.

Lueker, T.J, A.G. Dickson and C.D. Keeling, 2000. Ocean pCO2 calculated from dissolved inorganic carbon, alkalinity, and equations for K1 and K2: validation based on laboratory measurements of CO2 in gas and seawater at equilibrium. Marine Chem., 70, 105-119.

Saito, S., M. Ishii, T. Midorikawa and H.Y. Inoue 2008. Precise Spectrophotometric Measurement of Seawater pHT with an Automated Apparatus using a Flow Cell in a Closed Circuit. Technical Reports of Meteorological Research Institute, 57, 1-28.

- Chlorofluorocarbon (CFC-11 and CFC-12)

13 December 2013

(1) Personnel

Kazuki ISHIMARU (GEMD/JMA)

Etsuro ONO (GEMD/JMA)

(2)  Station occupied

A total of 27 stations (Leg 1: 15, Leg 2: 12) were occupied for analyses of CFCs (CFC-11 and CFC-12). Station location and sampling layers for CFCs analyses are shown in Figure C.7.1.

Figure C.7.1. Station location (left panel) and sampling layers of CFCs (right panel).

(3) Sampling

On the casts where water samples were collected for dissolved CFC-11, CFC-12 analysis, these were the first samples drawn from bottles to reduce CFC contamination. To minimize contact with air, the CFC samples were drawn directly through the stopcocks of the 10-liter Niskin bottles into 100 ml precision glass syringes equipped with three-way stopcocks. The syringes were immersed in a holding tank of clean surface seawater with ice before being analyzed.

(4) Instruments and methods

Concentrations of CFC-11 and CFC-12 in seawater and air samples were measured by shipboard gas chromatography (GC) equipped with an electron capture detector (ECD) using techniques described by Bullister and Weiss (1988). The flow diagram is shown in Figure C.7.2. For seawater analyses, sample water was transferred from a glass syringe to a fixed volume chamber (volume approximately 40 ml). The contents of the chamber were then injected into a purging chamber made of glass. The dissolved gases in the seawater sample were extracted by passing CFC-free purge gas through the purging chamber for a period of 6 minutes at 75 ml min-1. Water vapour was removed from the sample gas during passage through a 10 cm long, 3/8 inch outer diameter stainless tube packed with the desiccant magnesium perchlorate. The sample gases were concentrated on a cold-trap consisting of a 1/8 inch outer diameter stainless steel tube with a 5 cm section packed with Porasil C (80-100 mesh) and a 5 cm section packed with Porapak T (80-100 mesh). A cooler was used to cool the trap to approximately －35 °C. After 6 minutes of purging, the trap was isolated, and was transferred to the bath heated to approximately 95 °C. The sample gases held in the trap were then injected onto a precolumn (approximately 25 cm of 1/8 inch outer diameter stainless steel tubing packed with 80-100 mesh Porasil C, held at 70 °C) for the initial separation of CFC-12 and CFC-11. The CFCs that had passed from the pre-column entered the main analytical column (approximately 250 cm of 1/8 inch outer diameter stainless steel tubing packed with Porasil C, 80-100 mesh, held at 70 °C) of GC (a Shimadzu GC-8A with ECD).

Figure C.7.2. Schematic flow diagram of CFCs analytical system.

The analytical systems were calibrated using a standard gas of known CFC composition after analyses of a station. Gas sample loops of known volume were thoroughly flushed with standard gas and then the standard gas was injected into the system. Concentrations of CFCs standard gas were determined by a gravimetric technique developed by TAIYO NIPPON SANSO Corporation, Japan (Table C.7.1). The temperature and pressure was recorded so that the amount of gas injected could be calculated.

The procedures used to transfer the standard gas to the trap, precolumn, main chromatographic column, and ECD were similar to those used for analyzing water samples. Air samples and blank samples (CFC-free gas) were injected and analyzed in a similar manner with standard gases.

CFC concentrations in seawater and air samples were determined by the multi-point calibration curves made by fitting chromatographic peak areas in a quadratic polynomial. Two sizes of gas sample loops (2 ml and 10 ml) were prepared. Multiple injections of these loops volumes could be made to allow the system to be calibrated over a relatively wide range (from 0 to the concentration of air sample). The typical analysis time for seawater, air, standard gases or blank samples was approximately10 minutes.

Table C.7.1. Concentrations of standard gases.

| Cylinder No. | CFC-11 (pptv) | CFC-12 (pptv) |
| --- | --- | --- |
| CPB23436 | 50.20 | 49.35 |
| CPB29414 | 150.2 | 100.0 |
| CPB30485 | 349.3 | 200.0 |

(5) Results

CFCs concentrations in the air are reported as mole fraction CFC in dry gas, and are typically in the parts per trillion (ppt). Dissolved CFC concentrations in seawater are given in units of picomoles per kilogram seawater (pmol kg-1). Full-range calibration curves (quadratic fitting of chromatographic peak areas for the number of moles of CFCs injected) were made after every station analyses. Figure C.7.3 shows time-series of concentrations determined by self-calibrations among these cylinders. It suggested that the stability of CFCs cylinders was ensured during the cruise.

The repeatability for the standard gas in the fixed volume were 0.98% for CFC-12 and 0.92% for CFC-11 (n = 10). We estimated the detection limit for our analytical system to be 0.05 pmol kg-1 for both CFC-11 and CFC-12.

Figure C.7.3. Self-calibration of standard gas cylinder (left panel is CPB23436, centre panel is CPB29414, right panel is CPB30485). CFC-11 (top panel) and CFC-12 (bottom panel). Thick lines denote concentration of standard gases cylinders, which is assigned by gravimetric technique. Dash-dot lines denote average of self-calibration. Dash lines were its standard deviation.

The purging efficiency was estimated periodically by re-purging high concentration surface water samples and measuring this residual signal. At a flow rate of 75 ml min-1 for 6 minutes, purging efficiency were >99.9% for CFC-11 and CFC-12. And no correction for this has been applied to the reported water concentration values.

On this expedition, based on the analysis of 37 pairs of replicate for CFC-11 and 36 pairs of replicate for CFC-12, the average of differences between pairs were calculated to be 0.019 pmol kg-1 for CFC-11 and 0.020 pmol kg-1 for CFC-12 (Figure C.7.3). And, based on the analysis of 23 pairs of duplicate samples for CFC-11 and 20 pairs of duplicate samples for CFC-12, the average of differences between pairs were calculated to be 0.022 pmol kg-1 for CFC-11 and 0.022 pmol kg-1 for CFC-12 (Figure C.7.4).

Figure C.7.4. Result of replicate samplings during this cruise versus station number (top panel), sampling depth in dbar (second from the top panel) and concentration of CFCs (second from bottom panel). Bottom panel shows histogram of the result of replicate samplings. Left panel shows CFC-11 and right panel shows CFC-12.

Figure C.7.5. Same as Fig.C.7.3 but for duplicate samplings

Air measurement at each station was useful in order to check CFCs measurements by calculating CFC saturation levels in the surface water. Averages of atmospheric CFC-11 and CFC-12 concentrations were estimated to be 241.7 ± 5.1 ppt and 533.8 ± 7.1 ppt respectively, based on 49 of measurement of air samples (Figure C.7.5).

Figure C.7.6. Result of CFCs concentrations in the air versus station number. CFC-11 (left panel) and CFC-12 (right panel).

We drew pCFCs property-property plots with the atmospheric CFC history for flagging (Figure C.7.7).

Figure C.7.7. pCFCs property-property plots. Plus and circle show pCFCs in seawater samples and historical CFCs of air concentration, respectively.

A small number of water samples showed anomalously high CFC concentrations compared with the samples obtained at adjacent layers. These results occurred sporadically during the cruise and were not clearly associated with other features in the water column (e.g., anomalous dissolved oxygen, salinity, or temperature features). This suggests that these samples were probably contaminated with CFCs during the sampling or analysis processes.

Measured concentrations for these anomalous samples are included in the preliminary data, but are given a quality flag values of either 3 (questionable measurement) or 4 (bad measurement). A quality flag of 5 was assigned to the data for samples which were drawn from the Niskin bottle but never analyzed due to a variety of reasons (e.g., leaking stopcock, plunger jammed in syringe barrel). A total of 29 analyses of CFC-11 and 32 analyses of CFC-12 were assigned a quality flag of 3. A total of 47 analyses of CFC-11 and 41 analyses of CFC-12 were assigned a quality flag of 4.

Table C.7.2. Summary of assigned quality control flags

| Flag | Definition | CFC-11 | CFC-12 |
| --- | --- | --- | --- |
| 2 | Good | 658 | 662 |
| 3 | Questionable | 29 | 32 |
| 4 | Bad (Faulty) | 47 | 41 |
| 5 | Not reported | 7 | 7 |
| 6 | Mean of replicate measurements | 35 | 34 |
| Total number of samples |  | 776 | 776 |

We show the vertical section of CFC-11 and CFC-12 along WHP-P9 line (Figure C.7.8) observed in this work.

Figure C.7.8. Vertical section of CFC-11 and CFC-12 concentration (in pmol kg-1) along WHP-P9 line observed in this work. Blue dots indicate sampling layer.

Also we show the previous CFC-11 section along WHP-P9 line conducted by JMA in 1994 (Figure C.7.9), and vertical profiles of CFCs at cross-point with WHP-P2 conducted by Scripps Institution of Oceanography (SIO) in 2004 (Figure C.7.10) and WHP-P3 conducted by SIO in 1985 and by Japan Agency for Marine-Earth Science and Technology (JAMSTEC) in 2005/06 (C.7.11). The implication of these results will be discussed elsewhere.

Figure C.7.9. Section of CFC-11 concentration (in pmol kg-1) along WHP-P9 line observed in 1994 by JMA. Blue dots indicate sampling layer.

Figure C.7.10. Comparison of dissolved CFCs profiles at cross-point of (a) CFC-11 (left panel) and (b) CFC-12 (right panel). Circle and square show WHP-P9 revisit in 2010 by JMA and WHP-P2 in 2004 by SIO, respectively.

Figure C.7.11. Comparison of dissolved CFCs profiles at cross-point of (a) CFC-11 (left panel) and (b) CFC-12 (right panel). Triangle, square, circle show the WHP-P3 in 1985 by SIO, WHP-P3 in 2005/06 by JAMSTEC and WHP-P9 revisit in 2010 by JMA, respectively.

References

Bullister, J.L., and R.F. Weiss (1988): Determination of CC13F and CC12F2 seawater and air. Deep-Sea Res., 25, 839-853.

- Phytopigment (chlorophyll-a and phaeopigments)

13 December 2013

(1)  Personnel

Yusuke TAKATANI (GEMD/JMA)

Shinichiro UMEDA (GEMD/JMA)

(2)  Station occupied

A total of 50 stations (Leg 1: 29, Leg 2: 21) were occupied for phytopigment. Station location and sampling layers of phytopigment are shown in Figure C.8.1.

Figure C.8.1. Sation location (left panel) and sampling layers of phytopigment (right panels).

(3)  Reagents

N,N-dimethylformamide (DMF)

0.5 N hydrochloric acid (0.5N HCl)

Chlorophyll-a standard from Anacystis nidulans algae (Lot. BCBB4166) manufactured by Sigma Chemical Co.

Rhodamine WT manufactured by Turner Designs.

(4)  Instruments

Fluorometer; 10-AU (S/N:6718) manufactured by Turner Designs

Spectrophotometer; UV-1800 (S/N:A114547) manufactured by Shimadzu Co. Ltd.

Glass Fiber Filiter; Whatman GF/F filter (25 mm)

- Standardization

A chlorophyll-a standard calibration for fluorometric determination was performed by the method described by UNESCO (1994). Before standardization, fluorometer was calibrated by using 100 % DMF and a Rhodamine solution diluted to 1ppm with deionized water. Chlorophyll-a standard was dissolved in DMF. The concentration of chlorophyll-a solution was determined spectrophotometrically as follows;

Chl  concentration μgml =Achlspecific absorption coefficient

where Achl is the difference between absorbance at 663.8 nm and 750 nm. The specific absorption coefficient is 88.74 L/g·cm (Porra et al., 1989). Using this precise chlorophyll-a concentration, the linear calibration factor (fph) and the acidification coefficient (R) were calculated. fph was calibrated for each cuvette as the slope of the unacidified fluorometric reading vs. chlorophyll-a concentration calculated spectrophotometrically. R was calculated by averaging the ratio of the unacidified and acidified readings of pure chlorophyll-a. Table C.8.1 shows fph and R in this cruise.

Table C.8.1. fph and R determined by the standardization.

| Linear calibration factor (fph) | 5.13 |
| --- | --- |
| Acidification coefficient (R) | 1.848 |

(6)  Seawater sampling and measurement

Seawater samples were collected from 10-liters Niskin bottle attached the CTD-system and a stainless steel bucket for the surface in 200 ml. The seawater samples were immediately filtered through 25 mm GF/F filter by low vacuum pressure, and the particulate matter was made to adsorb to the filter. The filter was put into the vial containing 9 ml of DMF, then stored to extract phytogigment in the refrigerator for more than 24 hours at –30 deg-C until analysis.

After the extracts were put on the room temperature for at least one hour in the dark, only the extracts except the filter were decanted from the vial to the cuvette. Fluorometer readings for each cuvettes were taken before and after acidification with 1-2 drops 0.5 N HCl. Chlorophyll-a (Chl) and phaeopigment (Phaeo) concentration in the sample are calculated using the following equations;

Chl μgl = F0-Fafph∙R-1∙vV

Phaeo μgl = R∙F0-Fafph∙R-1∙vV

F0	= reading before acidification

Fa	= reading after acidification

R	= acidification coefficient (F0/Fa) for pure chlorophyll-a

fph	= linear calibration factor

v	= extraction volume

V	= sample volume

(7)  Quality control flag assignment

Quality flag values were assigned to phytopigment measurements using the code defined in IOCCP Report No.14 (Swift, 2010). Measurement flags of 2 (good), 3 (questionable), and 4 (bad) have been assigned (Table C.8.2).

Table C.8.2. Summary of assigned quality control flags.

| Flag | Definition | Chl | Phaeo |
| --- | --- | --- | --- |
| 2 | Good | 437 | 437 |
| 3 | Questionable | 0 | 0 |
| 4 | Bad (Faulty) | 12 | 12 |
| Total number |  | 449 | 449 |

References

Porra, R. J., W. A. Thompson and P. E. Kriedemann (1989): Determination of accurate coefficients and simultaneous equations for assaying chlorophylls a and b extracted with four different solvents: verification of the concentration of chlorophyll standards by atomic absotption spectroscopy. Biochem. Biophy. Acta, 975, 384-394

Swift, J. H. (2010): Reference-quality water sample data: Notes on acquisition, record keeping, and evaluation. IOCCP Report No.14, ICPO Pub. 134, 2010 ver.1

UNESCO (1994), Protocols for the joint global ocean flux study (JGOFS) core measurements: Measurement of chlorophyll a and phaeopigments by fluorometric analysis, IOC manuals and guides 29, Chapter 14.

- Lowered Acoustic Doppler Current Profiler

13 December 2013

(1)  Personnel

Tetsuya NAKAMURA (GEMD/JMA)

Yoshikazu HIGASHI (GEMD/JMA)

Tomoyuki KITAMURA (GEMD/JMA)

Takahiro SEGAWA (GEMD/JMA)

Keizo SHUTTA (GEMD/JMA)

Yasuaki BUNGI (GEMD/JMA)

(2)  Instrument and measurement

Direct flow measurement from sea surface to the bottom was carried out using a Lowered Acoustic Doppler Current Profiler (LADCP). The instrument used was the RDI Workhorse Monitor 307.2 kHz unit (S/N 13666; Teledyne RD Instruments, USA). The instrument was attached on the CTD frame, orientating downward. The CPU firmware version was 50.36.

One ping raw data were recorded. Settings for the collecting data were as listed in Table C.9.1. A total of 124 operations were made with the CTD observations. The performance of the LADCP was good between Stn.1 (RF3649) and Stn.42 (RF3690). From Stn.43 (RF3691) the echo intensity of beam 4 got weak, and from Stn.50 (RF3698) it might be broken down. From Stn.109 (RF3757), the beam 1 might be also broken down. And besides, data transfer errors often occurred during download process from the LADCP to the PC. So the data processing was performed for 101 stations.

Table C9.1. Setting for the correcting data.

| Bin length | 8m |
| --- | --- |
| Bin number | 25 |
| Error Threshold | 2000mm/s |
| Ping interval | 1.0sec |

(3)  Data process and result

Vertical profiles of velocity are obtained by the inversion method (Visbeck, 2002). Both the up and down casts are used for the inversion. Since the first bin from LADCP is influenced by the turbulence generated by CTD frame, the weight for the inversion is set to small value of 0.1. The GPS navigation data are used in the calculation of the reference velocities and the bottom-track data are used for the correction of the reference velocities. Shipboard ADCP (SADCP) data averaged for 5 minutes are also included in the calculation. The CTD data are used for the sound speed and depth calculation. IGRF (International Geomagnetic Reference Field) 11th generation data are used for calculating magnetic deviation to correct the direction of velocity. In the processing, we use Matlab routines (version 8b: 5 April 2004) provided by M. Visbeck and G. Krahmann. We set the weight for SADCP data in the calculation to 3.0, so vertical profiles of velocity obtained by the inversion method is similar to SADCP upper 1000 dbar. The uncertainty of velocity observed by SADCP is about 10 cm/s. So we regard the error velocity from LADCP upper 1000 dbar as about 10 cm/s. Figure C.8.1 and C.8.2 show the results of the zonal velocity (eastward is positive) and the meridional velocity (northward is positive), respectively. The major currents in the western Pacific such as the Kuroshio (34°N to 32°N), the Equatorial Under Current (EQ to 3°N), and New Guinea Coastal Under Current (around 2°S) appeared in the figures. Figure C.8.3 shows error velocity estimated by the inversion method. The error velocities are very small (less than 5 cm/s) upper 1000 dbar and adjacent to the bottom from Stn.1 to Stn.42. After Stn.43, the error velocity below 1000 dbar become larger and exceed 50 cm/s at maximum. This is because the echo intensity of beam 4 got weak down after Stn.43.

Figure C.9.1. The cross-section of zonal velocity (m/s, eastward is positive). Black line shows the Stn.43. The data south of the Stn.43 is doubtful due to malfunction of the instrument.

Figure C.9.2 The cross-section of meridional velocity (m/s, northward is positive). Black line shows the Stn.43. Note that southern of Stn.43 is doubtful.

Figure C.9.3. Cross-section of error velocity (m/s) estimated by the inversion method. Black line shows the Stn.43. Note that southern of Stn.43 is doubtful.

Reference

Visbeck, M. (2002): Deep velocity profiling using Lowered Acoustic Doppler Current Profilers: Bottom track and inverse solutions. J. Atmos. Oceanic Technol., 19, 794-807.