CHLOROFLUOROCARBON (CFC) AND SULFUR HEXAFLUORIDE (SF6) MEASUREMENTS ON THE 2007/2008 CLIVAR P18 EXPEDITION

PIs:            John Bullister and Mark Warner

Analysts: Leg 1  David Wisegarver and Robert Letscher

Leg 2  Nathaniel Nutter and Nick Beaird

(Prepared 2 May

2011)

A PMEL analytical system (Bullister and Wisegarver, 2008) was used for CFC-11, CFC-12 and sulfur hexafluoride (SF6) analyses on the 2007/2008 CLIVAR P18 expedition.  Samples for the analyses of dissolved CFC-11, CFC-12 and SF6 (‘CFC/SF6’) were drawn from approximately 3160 of the 5700 water samples collected during the expedition. Analysis of CFC-11, CFC-12, and SF6 were made on the same water sample. Measurements of carbon tetrachloride (CCl4) were made on a subset of these samples, primarily in cold, deepwater samples where non-conservative behavior of this compound is less significant. The analytical system was not optimized for CCl4 measurements on this cruise.  CCl4 measurements in samples shallower than 2000 m are not reported.  CCl4 measurements in samples collected at depths greater than 2000 m have been flagged as ‘3’ (questionable).

In general, the analytical system performed well on the cruise.  SF6 measurements were not made on a routine basis as part of WOCE or on earlier CLIVAR repeat hydrography cruises, and the SF6 work on CLIVAR P18 is considered to be a pilot study to demonstrate the feasibility of making these measurements more routinely on future CLIVAR cruises.  Typical dissolved SF6 concentrations in modern surface water are ~1-2 fmol kg-1 seawater (1 fmol= femtomole = 10-15 moles), approximately 1000 times lower than dissolved CFC-11 and CFC-12 concentrations.  The limits of detection for SF6 on CLIVAR P18 were approximately 0.02 fmol kg-1, giving a concentration range of ~ 100:1 for SF6 measurements at present in the ocean.  SF6 measurements in seawater remain extremely challenging. Improvements in the analytical sensitivity to this compound at low concentrations are essential to make these measurements more routine on future CLIVAR cruises

Water samples on CLIVAR P18 were collected in bottles designed at PMEL that use a modified end-cap design to minimize the contact of the water sample with the end-cap O-rings after closing. Two sizes of these bottles (11 or 12 liter) were used on the expedition. Stainless steel springs covered with a nylon powder coat were substituted for the internal elastic tubing provided with standard Niskin bottles. On the casts where water samples were collected for dissolved CFC-11, CFC-12 and SF6 analysis, these were the first samples drawn from the bottles. Care was taken to coordinate the sampling of CFC/SF6 with other samples to minimize the time between the initial opening of each bottle and the completion of sample drawing. Samples more easily impacted by gas exchange (dissolved oxygen, 3He, DIC and pH) were collected within several minutes of the initial opening of each bottle. To minimize contact with air, the CFC/SF6 samples were drawn directly through the stopcocks of the bottles into 250 ml precision glass syringes equipped with three-way plastic stopcocks. The syringes were immersed in a holding tank of clean surface seawater held at ~5oC until ~30 minutes before being analyzed.  At that time, the syringe was placed in a bath of surface seawater heated to ~30oC.

For atmospheric sampling, a ~75 m length of 3/8" OD Dekaron tubing was run from the CFC van located on the fantail to the bow of the ship. A flow of air was drawn through this line into the main laboratory using an Air Cadet pump. The air was compressed in the pump, with the downstream pressure held at ~1.5 atm. using a backpressure regulator. A tee allowed a flow of ~100 ml min-1 of the compressed air to be directed to the gas sample valves of the CFC/SF6 analytical systems, while the bulk flow of the air (>7 l min-1) was vented through the back-pressure regulator. Air samples were analyzed only when the relative wind direction was within 60 degrees of the bow of the ship to reduce the possibility of shipboard contamination.  Analysis of bow air was performed at ~25 locations along the cruise track. At each location, at least five air measurements were made to increase the precision of the measurements. The measured CFC-11 and CFC-12 and SF6 air concentrations are reported in Table 1.

Concentrations of CFC-11, CFC-12 and SF6 in air samples, seawater, and gas standards were measured by shipboard electron capture gas chromatography (EC-GC) using techniques modified from those described by Bullister and Weiss (1988) and Bullister and Wisegarver (2008), as outlined below.  For seawater analyses, water was transferred from a glass syringe to a glass-sparging chamber (volume ~190 ml). The dissolved gases in the seawater sample were extracted by passing a supply of CFC/SF6 free purge gas through the sparging chamber for a period of 6 minutes at ~175 ml min-1. Water vapor was removed from the purge gas during passage through an 18 cm long, 3/8" diameter glass tube packed with the desiccant magnesium perchlorate. The sample gases were concentrated on a cold-trap consisting of a 1/16" OD stainless steel tube with a 5 cm section packed tightly with Porapak Q (60-80 mesh) and a 22 cm section packed with Carbosieve G. A Neslab Cryocool CC-100 was used to cool the trap to ~70°C.  After 6 minutes of purging, the trap was isolated, and it was heated electrically to ~175°C. The sample gases held in the trap were then injected onto a precolumn (~60 cm of 1/8" O.D. stainless steel tubing packed with 80-100 mesh Porasil B, held at 80°C) for the initial separation of CFC-12, CFC-11, SF6  and CCl4 from later eluting peaks.

After the SF6 and CFC-12 had passed from the pre-column and into the second precolumn (5 cm of 1/8" O.D. stainless steel tubing packed with MS5A, 95°C) and into the analytical column #1 (240 cm of 1/8" OD stainless steel tubing packed with MS5A and held at 95°C), the outflow from the first precolumn was diverted to the second analytical column (150 cm 1/8" OD stainless steel tubing packed with Carbograph 1AC, 80-100 mesh, held at 80°C). After CFC-11 had passed through the first pre-column, the flow was diverted to a third analytical column (1 m , Carbograph 1AC, 80°C). The gases remaining after CCl4 had passed through the first pre-column, were backflushed from the pre column and vented. Column #1 and the second pre-column were held in a Shimadzu GC8 gas chromatograph with an electron capture detector (ECD) held at 340oC. Column #2 and #3, and the first precolumn were in another Shimadzu GC8 gas chromatograph with ECD. The outflow from column #3 was plumbed to a Shimadzu Mini2 gas chromatograph with the ECD held at 250oC.

The analytical system was calibrated frequently using a standard gas of known CFC/SF6 composition. Gas sample loops of known volume were thoroughly flushed with standard gas and injected into the system.  The temperature and pressure was recorded so that the amount of gas injected could be calculated. The procedures used to transfer the standard gas to the trap, precolumn, main chromatographic column, and ECD were similar to those used for analyzing water samples. Four sizes of gas sample loops were used. Multiple injections of these loop volumes could be made to allow the system to be calibrated over a relatively wide range of concentrations. Air samples and system blanks (injections of loops of CFC/SF6 free gas) were injected and analyzed in a similar manner. The typical analysis time for seawater, air, standard or blank samples was ~11 minutes.  Concentrations of the CFC-11 and CFC-12 in air, seawater samples, and gas standards are reported relative to the SIO98 calibration scale (Cunnold et al., 2000). Concentrations of SF6 in air, seawater samples, and gas standards are reported relative to the CMDL calibration scale (Bullister et al, 2006). Concentrations in air and standard gas are reported in units of mole fraction CFC in dry gas, and are typically in the parts per trillion (ppt) range. Dissolved CFC concentrations are given in units of picomoles per kilogram seawater (pmol kg-1) and SF6 concentrations in fmol kg-1.  CFC/SF6 concentrations in air and seawater samples were determined by fitting their chromatographic peak areas to multi-point calibration curves, generated by injecting multiple sample loops of gas from a working standard (PMEL cylinder 45186) into the analytical instrument. The response of the detector to the range of moles of CFC/SF6 passing through the detector remained relatively constant during the cruise. Full-range calibration curves were run at intervals of 4-5 days during the cruise. Single injections of a fixed volume of standard gas at one atmosphere were run much more frequently (at intervals of ~90 minutes) to monitor short-term changes in detector sensitivity.

The purging efficiency was estimated by re-purging a high-concentration water sample and measuring this residual signal.  At a flow rate of 170 cc min-1 for 6 minutes, the purging efficiency for all 3 gases was >99.9% and no correction for this has been applied to the reported water concentration values.

On this expedition, based on the analysis of ~250 pairs of duplicate samples, we estimate precisions (1 standard deviation) of about 1% or 0.002 pmol kg-1 (whichever is greater) for both dissolved CFC-11 and CFC-12  measurements.  The estimated precision for SF6 was 2% or 0.02 fmol kg-1, (whichever is greater). Overall accuracy of the measurements (a function of the absolute accuracy of the calibration gases, volumetric calibrations of the sample gas loops and purge chamber, errors in fits to the calibration curves and other factors) is estimated to be about 2% or 0.004 pmol kg-1 for CFC11 and CFC-12 and 4% or 0.04 fmol kg-1 for SF6). As discussed above, carbon tetrachloride was analyzed as a qualitative tracer only on this expedition and all reported concentration values (from samples deeper than 2000 m) have been flagged as ‘3’ (questionable).

Based on the earlier occupation of the P18 section in 1994 as part of the WOCE program and crossings of this section by other WOCE-era sections, the region of the water column between 5oN and 20oS and deeper than 2000 m was thought have near zero levels of CFCs and SF6 at the time of the CLIVAR P18 section.  The means of measured values in this region on the CLIVAR P18 expedition was ~0.0025 pmol kg-1 for CFC-11, 0.001 pmol kg-1 for CFC-12 and 0 fmol kg-1 for SF6.  Based on previous experiments showing a slow grow-in of CFCs in water held in the sample bottles, and the scatter in measured concentrations of these deep samples, we estimate a total sampling and analytical blank of 0.0025 pmol kg-1 for CFC-11, 0.001 pmol kg-1 for CFC-12 and 0 fmol kg-1 for SF6. The final water concentration data reported here have had these blank corrections applied.

A small number of water samples had anomalously high CFC/SF6 concentrations relative to adjacent samples. These samples occurred sporadically during the cruise and were not clearly associated with other features in the water column (e.g., anomalous dissolved oxygen, salinity, or temperature features). This suggests that these samples were probably contaminated with CFCs/SF6 during the sampling or analysis processes.

Measured concentrations for these anomalous samples are included in the data file, but are given a quality flag value of either 3 (questionable measurement) or 4 (bad measurement). Approximately 68 samples out of 3168 (~1.7%) CFC-11, CFC-12 and SF6 samples were assigned quality flags of 3 or 4.  A quality flag of 5 was assigned to samples which were drawn from the rosette but never analyzed due to problems during handling or analysis (e.g., leaking, broken syringe etc.).

References

Bullister, J.L., and R.F. Weiss, 1988: Determination of CC13F and CC12F2 in seawater and air. Deep-Sea Res., v. 25, pp. 839-853.

Bullister, J.L., and D.P. Wisegarver (2008): The shipboard analysis of trace levels of sulfur hexafluoride, chlorofluorocarbon-11 and chlorofluorocarbon-12 in seawater. Deep-Sea Res. I, 55, 1063–1074.

Bullister, J.L., D.P. Wisegarver, and R.E. Sonnerup (2006): Sulfur hexafluoride as a transient tracer in the north Pacific Ocean. Geophys. Res. Lett., 33, L18603, doi: 10.1029/2006GL026514

Prinn, R.G., R.F. Weiss, P.J. Fraser, P.G. Simmonds, D.M.  Cunnold, F.N. Alyea, S. O'Doherty, P. Salameh, B.R.  Miller, J. Huang, R.H.J. Wang, D.E. Hartley, C. Harth,  L.P. Steele, G. Sturrock, P.M. Midgley, and A. McCulloch, 2000: A history of chemically and radiatively important gases in air deduced from ALE/GAGE/AGAGE. J. Geophys. Res., v. 105, pp. 17,751-17,792.

Table 1. Air Measurements from 2007/2008 CLIVAR P18

Date   Time     Lat     Lon     CFC12  CFC11"   SF6

YYMMDD HHMM

071217  0311     27.9   -112.9  539.2           6.30

071219  0255     20.8   -110.0  536.7   243.7   6.12

071222  0034     13.5   -110.0  536.2   243.7   6.58

071225  0009      8.0   -110.0  540.4           6.15

071226  0022      6.2   -110.0  540.8   243.6   6.25

071227  2242      2.0   -110.0  536.1   242.9   6.14

071229  0038      1.5   -110.0  538.1   242.9   6.12

080101  0607     -3.1   -110.0  537.7   243.8   5.88

080104  0958     -6.3   -108.3  539.1   243.9   6.02

080108  0921    -12.3   -103.0  533.6   243.0   6.12

080114  1018    -23.4   -103.0  537.4   242.3   5.98

080116  1943    -27.6   -103.0  529.4   242.1   6.06

080118  0740    -28.0   -103.0  528.8   242.9   5.78

080121  0919    -27.3   -109.0  543.9   241.9   5.96

080125  0503    -32.7   -103.0  534.6   240.1   6.02

080128  0426    -39.2   -103.0  534.0   239.9   6.04

080206  0326    -55.5   -103.0  537.1   239.4   5.93

080206  1912    -55.5   -103.0  531.5   238.9   6.03

080208  0041    -59.3   -103.0  532.9   238.5   6.12

080208  1545    -59.8   -103.0  532.2   232.1   5.99

080209  0525    -60.8   -103.0  535.9   238.6   6.02

080210  2224    -63.7   -103.0  531.4   237.3   5.88

080211  2216    -65.2   -103.0  532.5   240.2   5.95

080214  2153    -69.4   -103.5  533.2   239.7   5.88

080216  1336    -67.0   -107.3  533.2   242.6   5.90