Chlorofluorocarbon (CFC) and Sulfur Hexafluoride (SF6) measurements on CLIVAR Cruise A10 in 2011

PI:            John Bullister

Analysts: 	David Wisegarver

Alan Foreman

Darren Pilcher

A PMEL analytical system (Bullister and Wisegarver, 2008) was used for CFC-11, CFC-12 and sulfur hexafluoride (SF6) analyses on the CLIVAR A10 expedition.  Approximately 2300 samples of dissolved CFC-11, CFC-12 and SF6 (‘CFC/SF6’) were analyzed. The system was modified to include analyses of dissolved carbon tetrachloride (CCl4) and nitrous oxide (N2O).  The nitrous oxide measurements are considered exploratory and are not included in this report.  The carbon tetrachloride measurements are also considered exploratory, but are included in the data report. Because of analytical problems, all of the reported CCl4 concentrations are given either WOCE quality flags of ‘3’ (questionable) or ‘4’ (bad).

In general, the analytical system for CFC-11, CFC-12 and SF6 performed well on the cruise.  Typical dissolved SF6 concentrations in modern surface water are ~1-2 fmol kg-1 seawater (1 fmol= femtomole = 10-15 moles), approximately 1000 times lower than dissolved CFC-11 and CFC-12 concentrations.  The limits of detection for SF6 were approximately 0.02 fmol kg-1.  SF6 measurements in seawater remain extremely challenging. Improvements in the analytical sensitivity to this compound at low concentrations are essential to make these measurements more routine on future CLIVAR cruises

Water samples were collected in bottles designed with a modified end-cap to minimize the contact of the water sample with the end-cap O-rings after closing.  Stainless steel springs covered with a nylon powder coat were substituted for the internal elastic tubing provided with standard Niskin bottles. When taken, water samples collected for dissolved CFC-11, CFC-12 and SF6 analysis were the first samples drawn from the bottles. Care was taken to coordinate the sampling of CFC/SF6 with other samples to minimize the time between the initial opening of each bottle and the completion of sample drawing. Samples easily impacted by gas exchange (dissolved oxygen, 3He, DIC and pH) were collected within several minutes of the initial opening of each bottle. To minimize contact with air, the CFC/SF6 samples were drawn directly through the stopcocks of the bottles into 250 ml precision glass syringes equipped with three-way plastic stopcocks. To minimize degassing and possible bubble formation, the syringes were immersed in a holding tank of clean surface seawater held at ~10OC until ~20 minutes before being analyzed.  At that time, the syringe was place in a bath of surface seawater heated to ~30oC.

For atmospheric sampling, a ~75 m length of 3/8" OD Dekaron tubing was run from the CFC van located on the fantail to the bow of the ship. A flow of air was drawn through this line into the main laboratory using an Air Cadet pump. The air was compressed in the pump, with the downstream pressure held at ~1.5 atm. using a backpressure regulator. A tee allowed a flow of ~100 ml min-1 of the compressed air to be directed to the gas sample valves of the CFC/SF6 analytical systems, while the bulk flow of the air (>7 l min-1) was vented through the back-pressure regulator. Air samples were analyzed only when the relative wind direction was within 60 degrees of the bow of the ship to reduce the possibility of shipboard contamination.  Analysis of bow air was performed at ~10 locations along the cruise track. At each location, at least five air measurements were made to increase the precision of the measurements

Concentrations of CFC-11, CFC-12 and SF6 in air samples, seawater, and gas standards were measured by shipboard electron capture gas chromatography (EC-GC) using techniques modified from those described by Bullister and Weiss (1988) and Bullister and Wisegarver (2008), as outlined below.  For seawater analyses, water was transferred from a glass syringe to a glass-sparging chamber (volume ~200 ml). The dissolved gases in the seawater sample were extracted by passing a supply of CFC/SF6 free purge gas through the sparging chamber for a period of 6 minutes at ~150 ml min-1. Water vapor was removed from the purge gas during passage through an 18 cm long, 3/8" diameter glass tube packed with the desiccant magnesium perchlorate. The sample gases were concentrated on a cold-trap consisting of a 1/16" OD stainless steel tube with a 2.5 cm section packed tightly with Porapak Q (60-80 mesh), a 15 cm section packed with Carboxen 1000 and a 2.5 cm section packed with MS5A. A Neslab Cryocool CC-100 was used to cool the trap to ~-65°C.  After 6 minutes of purging, the trap was isolated, and it was heated electrically to ~175°C.

Once the trap was hot, the gases were injected onto a series of precolumns and columns.  There were three main analytical columns, one for SF6 and CFC-12, one for CFC-11 and CCL4 and the last for N2O.  The injection began with all of the gases passing through the first precolumn (PC1, ~45 cm of 1/8" O.D. stainless steel tubing packed with 80-100 mesh Porasil B, held at 80°C).  The first precolumn (PC1) had several purposes.  The first was to separate N2O, SF6 and CFC-12 from CFC-11 so that these gases could pass on to the second precolumn (PC2) without any CFC-11.  Secondly, the PC1 separated CFC-11 and CCL4 from later peaks, so chromatograms could be kept moderately short.  Once CCL4 had passed through the PC1 and onto the analytical column(180 cm 1/8" OD stainless steel tubing packed with Porasil B, 80-100 mesh, held at 80°C), PC1 was backflushed to remove the unwanted peaks. While CFC-11 and CCL4 were passing through PC1, N2O, SF6 and CFC-12 were passing through the PC2 (5 cm of 1/8" O.D. stainless steel tubing packed with MS5A, 120°C).   SF6 and CFC-12 pass through PC2 relatively quickly leaving N2O which was then diverted to its analytical column (2 m , Hayesep B, 120°C, ECD temperature of 330° C) with a flow of Argon-Methane (95:5).

Subsequent to passing through PC2, SF6 and CFC-12 were separated on

an analytical column (180 cm 1/8" OD stainless steel tubing packed

with MS5A, 60-80 mesh plus 60 cm of 1/8" OD stainless steel tubing

packed with Porasil B, 80-100 mesh, held at 80°C, ECD temperature at

340° C).

The analytical column for N2O and the second pre-column were held in a Shimadzu GC8 gas chromatograph with an electron capture detector (ECD) held at 330oC. The other two columns and the first precolumn were in another Shimadzu GC8 gas chromatograph with ECD (340o C). The outflow from the column for CFC-11 and CCL4 was directed to a Shimadzu Mini2 gas chromatograph (no column) with the ECD held at 250oC.

The analytical system was calibrated frequently using a standard gas of known CFC/SF6 composition. Gas sample loops of known volume were thoroughly flushed with standard gas and injected into the system.  The temperature and pressure was recorded so that the amount of gas injected could be calculated. The procedures used to transfer the standard gas to the trap, pre-column, main chromatographic column, and ECD were similar to those used for analyzing water samples. Four sizes of gas sample loops were used. Multiple injections of these loop volumes could be made to allow the system to be calibrated over a relatively wide range of concentrations. Air samples and system blanks (injections of loops of CFC/SF6 free gas) were injected and analyzed in a similar manner. The typical analysis time for seawater, air, standard or blank samples was ~11 minutes.

Concentrations of CFC-11, CFC-12 and SF6 in air were calibrated relative to PMEL standard cylinder PMEL-72584. Concentrations of CFC-11, CFC-12 and SF6 in seawater were calibrated relative to PMEL standard cylinder PMEL-72611. CFC-11 and CFC-12 concentrations are reported on the SIO93 scale (Cunnold et al., 2000) and SF6 concentrations are reported relative SIO 2005 scale.

Concentrations in air and standard gas are reported in units of mole fraction CFC in dry gas, and are typically in the parts per trillion (ppt) range.

The mean air concentrations measured on the A10 cruise were:

CFC-11: 240.0 ppt (+-0.7%)

CFC-12: 524.1 ppt (+-1.1%)

SF6   :   7.2 ppt (+-2.5%)

Dissolved CFC concentrations are given in units of picomoles per kilogram seawater (pmol kg-1) and SF6 concentrations in fmol kg-1.  CFC/SF6 concentrations in air and seawater samples were determined by fitting their chromatographic peak areas to multi-point calibration curves, generated by injecting multiple sample loops of gas from a working standard (PMEL cylinder 72611) into the analytical instrument. The response of the detector to the range of moles of CFC/SF6 passing through the detector remained relatively constant during the cruise. Full-range calibration curves were run at intervals of 4-5 days during the cruise. Single injections of a fixed volume of standard gas at one atmosphere were run much more frequently (at intervals of ~90 minutes) to monitor short-term changes in detector sensitivity.

The purging efficiency was estimated by re-purging a high-concentration water sample and measuring this residual signal.  At a flow rate of 150 cc min-1 for 6 minutes, the purging efficiency for CFC/SF6 gases was > 98.9%.  The purging efficiency for N2O was ~88%.  Reported concentrations are corrected for these efficiencies.

On some oceanographic expeditions, estimates of seawater CFC sampling blanks can be made by analyzing samples collected in areas thought to be essentially free of CFCs at the time of sampling.  None of the deep water along the A10 section was thought to be completely free on CFCs in 2011.  Based on the low and relatively uniform concentrations of CFC-11 (~0.005 +-0.002 pmol kg-1), CFC-12 (~0.004 +- 0.002 pmol kg-1) and SF6 (~0 fmol kg-1) measured in seawater samples between 2500 m and 3500 m depth along the A10 section east of about 5oE, we believe the CFC-11 and CFC-12 sampling blanks to be < 0.002 pmol kg-1 and the SF6 sampling blank to be < 0.01 fmol kg-1

No corrections for these low level sampling blanks have been applied to the reported CFC-11, CFC-11 and SF6 seawater concentrations.

On this expedition, based on the analysis of more than 150 pairs of duplicate samples, we estimate precisions (1 standard deviation) of about 0.2% or 0.002 pmol kg-1 (whichever is greater) for both dissolved CFC-11 and CFC-12  measurements.  The estimated precision for SF6 was 2% or 0.02 fmol kg-1, (whichever is greater). Overall accuracy of the measurements (a function of the absolute accuracy of the calibration gases, volumetric calibrations of the sample gas loops and purge chamber, errors in fits to the calibration curves and other factors) is estimated to be about 2% or 0.004 pmol kg-1 for CFC11 and CFC-12 and 4% or 0.04 fmol kg-1 for SF6).

A small number of water samples had anomalously high CFC/SF6 concentrations relative to adjacent samples. In most cases*, these samples occurred sporadically during the cruise and were not clearly associated with other features in the water column (e.g., anomalous dissolved oxygen, salinity, or temperature features). This suggests that these samples were probably contaminated with CFCs/SF6 during the sampling or analysis processes.

Measured concentrations for these anomalous samples are included in the data file, but are given WOCE quality flag values of either 3 (questionable measurement) or 4 (bad measurement.

The following number of samples were given flags of 3:

CFC-11:	   6

CFC-12:	   6

SF6:		   9

CCl4:		1898

The following number of samples were given flags of 4:

CFC-11:	  18

CFC-12:	  19

SF6:		  15

CCl4:		4446

A quality flag of 5 was assigned to samples which were drawn from the rosette but never analyzed due to a variety of reasons (e.g., leaking stopcock, plunger jammed in syringe barrel, analytical malfunction, etc.).

* A large number of samples in the deep (>3000 m) Brazil Basin west of ~10oW had anomalously high SF6 concentrations relative to the CFC-11 and CFC-12 concentrations.  These high SF6 concentrations are thought to be due to earlier deliberate deep SF6 tracer release experiments in this region (Rye et al., 2012).

References

Bullister, J.L., and R.F. Weiss, 1988: Determination of  CC13F and CC12F2 in seawater and air. Deep-Sea Res., v. 25,  pp. 839-853.

Bullister, J.L., and D.P. Wisegarver (2008): The shipboard analysis of trace levels of sulfur hexafluoride, chlorofluorocarbon-11 and chlorofluorocarbon-12 in seawater. Deep-Sea Res. I, 55, 1063–1074.

Prinn, R.G., R.F. Weiss, P.J. Fraser, P.G. Simmonds, D.M.  Cunnold, F.N. Alyea, S. O'Doherty, P. Salameh, B.R.  Miller, J. Huang, R.H.J. Wang, D.E. Hartley, C. Harth,  L.P. Steele, G. Sturrock, P.M. Midgley, and A. McCulloch, 2000: A history of chemically and radiatively important gases in air deduced from ALE/GAGE/AGAGE. J. Geophys.  Res., v. 105, pp. 17,751-17,792.

Rye, C.D., Messias, M-J, Ledwell, J.R., Watson, A.J., Brousseau, A., and King, B.A., 2012: Diapycnal diffusivities from a tracer release experiment in the deep sea, integrated over 13 years.  GRL Vol 39, L04603.  Doi:10.1029/2011GL050294, 2012.