Update of bottle oxygen data

(November 26, 2013)

We have conducted the inter-laboratory comparison of the standard potassium iodate solution between JMA and Japan Agency for Marine-Earth Science and Technology (JAMSTEC) in order to ensure the comparability of dissolved oxygen concentration in sea water since 2010. Moreover, we also conducted the inter-laboratory comparison of the standard potassium iodate solution between JMA and Scripps Institution of Oceanography (SIO) on board R/V Melville in May 2013. Table 1 shows pure KIO3 reagents used in each laboratory.

The results of inter-laboratory comparisons are summarized as follows. The standard potassium iodate solution made by Merck reagent shows a systematic difference to the solutions made from Wako reagent and Alfa Aesar reagent. On the other hand, the standard potassium iodate solution made by National Metrology Institute of Japan (NMIJ) reagent has no significant difference to the standard solutions made from Wako reagent and Alfa Aesar reagent (Table 2 and Figure 1).

The measured normality of standard potassium iodate solution made by Merck reagent based on the nominal normality of NMIJ reagent is higher (0.26±0.12%) than the nominal normality (Figure 2). From t-test, this difference is significant. Thus, the normality of the standard potassium iodate solution made by Merck reagent should be corrected by the factor as 1.0026 from the result of the inter-laboratory comparison studies conducted by JMA.

Acknowledgements

We are grateful to Michio Aoyama of the Meteorological Research Institute for helpful advice and discussion. We also thank Yuichiro Kumamoto of JAMSTEC, Fuyuki Shibata of Marine Works Japan, James H. Swift and Susan M. Becker of SIO for providing their comparison results.

Table 1. List of pure KIO3 reagent used in each institution

| Institution | Company | Lot. | Purity (%) |
| --- | --- | --- | --- |
| JMA | Merck | 62404E | 99.75 (2010–2011) |
|  |  | 92404G | 99.74 (2011–) |
|  | NMIJ | CRM 3006-a | 99.973 (2012–) |
| JAMSTEC | Wako | EPR3227 | >99.95 |
| SIO | Alfa Aesar | B05N35 | 99.4 – 100.4 |

Table 2. The results of standard potassium iodate solution inter-comparison

| Date | measured (N) | Nominal (N) | Uncertainty* (N) | Ratio (measured/nominal) |
| --- | --- | --- | --- | --- |
| (1) normality of the standard potassium iodate solution made by Wako versus Merck at JMA |
| 2010/5/23 | 0.009985 | － | 0.000008 | － |
| 2010/5/29 | 0.010004 | － | 0.000007 | － |
| 2010/12/18 | 0.009983 | － | 0.000009 | － |
| 2011/6/22 | 0.009981 | 0.010014 | 0.000013 | 0.9967 |
| 2011/6/22 | 0.009978 | 0.010014 | 0.000013 | 0.9964 |
| 2011/6/22 | 0.009986 | 0.010014 | 0.000012 | 0.9972 |
| 2011/6/22 | 0.009981 | 0.010014 | 0.000012 | 0.9967 |
| 2011/6/22 | 0.009984 | 0.010014 | 0.000013 | 0.9970 |
| 2011/6/22 | 0.009988 | 0.010014 | 0.000012 | 0.9974 |
| 2012/5/28 | 0.009993 | 0.010009 | 0.000008 | 0.9984 |
| 2012/5/28 | 0.009986 | 0.010010 | 0.000009 | 0.9976 |
| 2012/5/28 | 0.009989 | 0.010012 | 0.000008 | 0.9977 |
| 2012/5/28 | 0.009986 | 0.010012 | 0.000009 | 0.9974 |
| 2012/5/28 | 0.009989 | 0.010012 | 0.000008 | 0.9977 |
| 2012/5/28 | 0.009990 | 0.010012 | 0.000008 | 0.9978 |
| 2013/6/10 | 0.009976 | 0.010002 | 0.000009 | 0.9974 |
| 2013/6/10 | 0.009979 | 0.010002 | 0.000009 | 0.9977 |
| 2013/6/10 | 0.009974 | 0.010002 | 0.000009 | 0.9972 |
| Mean ratio of the measured normality versus the nominal normality = 0.9974 ± 0.0005 |
| (2) normality of the standard potassium iodate solution made by Wako versus NMIJ at JMA |
| 2012/5/28 | 0.010015 | 0.010009 | 0.000008 | 1.0006 |
| 2012/5/28 | 0.010008 | 0.010010 | 0.000009 | 0.9998 |
| 2012/5/28 | 0.010012 | 0.010012 | 0.000008 | 1.0000 |
| 2012/5/28 | 0.010008 | 0.010012 | 0.000009 | 0.9996 |
| 2012/5/28 | 0.010011 | 0.010012 | 0.000008 | 0.9999 |
| 2012/5/28 | 0.010013 | 0.010012 | 0.000008 | 1.0001 |
| 2013/6/10 | 0.010001 | 0.010002 | 0.000009 | 0.9999 |
| 2013/6/10 | 0.010003 | 0.010002 | 0.000010 | 1.0001 |
| 2013/6/10 | 0.009999 | 0.010002 | 0.000009 | 0.9997 |
| Mean ratio of the measured normality versus the nominal normality = 1.0000 ± 0.0003 |
| *: The standard uncertainty estimated theoretically by precision of burette, etc. Table 2. (Continued). |
| Date | measured (N) | Nominal (N) | Uncertainty** (N) | Ratio (measured/nominal) |
| (3) normality of the standard potassium iodate solution made by Merck versus Wako at JAMSTEC |
| 2010/7/14 | 0.009976 | 0.009955 | 0.000003 | 1.0021 |
| 2010/7/14 | 0.009990 | 0.009972 | 0.000003 | 1.0018 |
| 2010/7/14 | 0.009990 | 0.009972 | 0.000003 | 1.0018 |
| 2010/7/14 | 0.010002 | 0.009977 | 0.000003 | 1.0025 |
| 2010/9/3 | 0.009994 | 0.009967 | 0.000003 | 1.0027 |
| 2010/9/3 | 0.009979 | 0.009952 | 0.000003 | 1.0027 |
| 2011/5/30 | 0.010250 | 0.010233 | 0.000003 | 1.0017 |
| 2011/5/30 | 0.010168 | 0.010154 | 0.000003 | 1.0014 |
| 2011/5/30 | 0.010017 | 0.010001 | 0.000003 | 1.0016 |
| 2011/5/31 | 0.010573 | 0.010546 | 0.000003 | 1.0026 |
| 2011/5/31 | 0.010571 | 0.010546 | 0.000003 | 1.0024 |
| 2011/5/31 | 0.010856 | 0.010840 | 0.000003 | 1.0015 |
| 2012/4/24 | 0.010168 | 0.010148 | 0.000003 | 1.0020 |
| 2012/4/24 | 0.010049 | 0.010031 | 0.000003 | 1.0018 |
| 2012/4/24 | 0.010050 | 0.010031 | 0.000003 | 1.0019 |
| 2012/4/25 | 0.010112 | 0.010094 | 0.000003 | 1.0018 |
| 2012/4/25 | 0.010189 | 0.010162 | 0.000003 | 1.0027 |
| 2012/4/25 | 0.010018 | 0.010001 | 0.000003 | 1.0017 |
| 2013/5/21 | 0.010010 | 0.009989 | 0.000003 | 1.0021 |
| 2013/5/21 | 0.010012 | 0.009989 | 0.000003 | 1.0023 |
| 2013/5/21 | 0.010006 | 0.009989 | 0.000003 | 1.0017 |
| 2013/5/21 | 0.010014 | 0.009990 | 0.000003 | 1.0024 |
| 2013/5/21 | 0.010011 | 0.009990 | 0.000003 | 1.0021 |
| 2013/5/21 | 0.010002 | 0.009990 | 0.000003 | 1.0012 |
| Mean ratio of the measured normality versus the nominal normality = 1.0020 ± 0.0004 |
| (4) normality of the standard potassium iodate solution made by NMIJ versus Wako at JAMSTEC |
| 2012/4/25 | 0.009997 | 0.009997 | 0.000001 | 1.0000 |
| 2013/5/19 | 0.009995 | 0.009994 | 0.000001 | 1.0001 |
| 2013/5/19 | 0.009994 | 0.009994 | 0.000001 | 1.0000 |
| 2013/5/19 | 0.009980 | 0.009994 | 0.000001 | 0.9986 |
| 2013/5/21 | 0.009981 | 0.009994 | 0.000001 | 0.9987 |
| Mean ratio of the measured normality versus the nominal normality = 0.9995 ± 0.0008 |
| (5) normality of the standard potassium iodate solution made by Merck versus Alfa Aesar at SIO |
| 2013/5/10 | 0.010023 | 0.009996 | 0.000003 | 1.0027 |
| 2013/5/10 | 0.010062 | 0.010033 | 0.000003 | 1.0029 |
| 2013/5/10 | 0.010050 | 0.010033 | 0.000003 | 1.0017 |
| Mean ratio of the measured normality versus the nominal normality = 1.0024 ± 0.0006 |
| **: The standard uncertainty by gravimetric method Table 2. (Continued). |
| (6) normality of the standard potassium iodate solution made by NMIJ versus Alfa Aesar at SIO |
| 2013/5/10 | 0.009994 | 0.009994 | 0.000001 | 1.0000 |
| 2013/5/10 | 0.009990 | 0.009994 | 0.000001 | 0.9996 |
| Mean ratio of the measured normality versus the nominal normality = 0.9998 ± 0.0003 |

Figure 1. Results of standard potassium iodate solution inter-laboratory comparison among JMA, JAMSTEC and SIO. Black (Gray) closed circles denote measured/nominal ratio of Wako (Merck) versus Merck (Wako), and black (gray) closed diamonds denote measured/nominal ratio of Wako (NMIJ) versus NMIJ (Wako). Gray open circles (diamonds) denote measured/nominal ratio of Merck (NMIJ) versus Alfa Aesar.

Figure 2. Results of inter-comparison between Merck and NMIJ. Upper panel shows measured/nominal ratio of Merck versus NMIJ for each standard potassium iodate solution batch serial, and lower panel shows the histogram of its ratios.