CO2/CLIVAR Repeat Hydrography Program 2011 Reoccupation of

WOCE Section A10

NOAA ship RONALD H. BROWN

Cape Town, South Africa – Rio de Janeiro, Brazil

September 26 – October 31, 2011

Chief Scientist: Molly Baringer, AOML

Co-Chief Scientist:  Alison MacDonald, WHOI

Data Manager:  Alex Quintero, Scripps

CTD Quality Control/Processing: Kristy McTaggart, PMEL

CTD Watchstanders: Elizabeth Simons, FSU and James Hooper, Scripps

Sample Salinity Analysts: Kyle Seaton, CIMAS and Andy Stefanick, AOML

Sample Oxygen Analysts: Chris Langdon, RSMAS and George Berberian, AOML

LADCP Technician: Elizabeth Simons, FSU for Sarah Eggleston, UH

Survey Technicians: Jonathan Shannahoff and Laurie Roy

Ship’s Electronics Technician: Jeff Hill

Summary

This cruise was a reoccupation of a longitudinal section nominally along 30S (WOCE Section A10, occupied in 2003 and 1993). Operations included CTDO/LADCP/rosette casts nominally at half-degree spacing.  Underway data collected included upper-ocean currents from the shipboard ADCP, surface oceanographic and meteorological parameters from the ship’s underway systems, and bathymetric data.  Ancillary operations included surface drifter deployments and Argo float deployments.

NOAA Ship Ronald H. Brown first departed Cape Town, South Africa on August 28, 2011 at 1000 UTC.  A successful test cast to 462 meters was completed on August 30.  After three days at sea, the ship returned to Cape Town on August 31 at 0700 UTC for repairs to the port thruster.  After a 12-day delay, the ship was underway again on September 13 at 1200 UTC for sea trials, which were unsuccessful, so the ship returned to Cape Town for another 12-day delay.  Finally, the ship departed on September 26, 2011 at 1000 UTC to begin the cruise after successful sea trials and a successful test cast to 400 meters.  The cruise ended in Rio de Janeiro, Brazil on October 31, 2011.

A total of 120 stations were occupied during A10 and 122 CTDO/LADCP/rosette casts were collected, including a designated black carbon cast at station 35 and a second cast at station 51 after a winch failure.  Fifteen Argo floats and ten surface drifters were deployed.  CTDO data, LADCP data, and water samples (up to 24) were collected on most casts, in most cases to within 10 meters of the bottom.

Salinity, dissolved oxygen, and nutrient samples were analyzed for up to 24 water samples from each cast of the principal CTDO/LADCP/rosette program. Water samples were also measured for CFCs, pCO2, Total CO2 (DIC), Total Alkalinity, and pH.  Additional samples were collected for 3He, Tritium, 13C/ 14C, DOC, DON, and POC.

CTD Underwater Package

Sea-Bird instrumentation was mounted in a 24-position aluminum frame provided by AOML with 24 10- and 11-liter Niskin bottles and AOML 24-position carousel s/n nnn.  Sea-Bird sensors initially on the 24-position frame included AOML’s 9plus CTD s/n 1035 and PMEL’s TCO sensors:  primary TCO s/n 03P-4569, 04C-3157, 43-0313 with 05T-1211; and secondary TCO s/n 03P-4341, 04C-2887, 43-1890 with 05T-5416.  Also mounted on the underwater package were an RDI Workhorse 150 kHz LACDP, Simrad altimeter, Wetlabs FLRTD fluorometer s/n 2125, and Benthos pinger s/n 1006.  A few changes were made to the underwater package as the cruise progressed, most notably the addition of an internally recording reference temperature sensor (SBE35RT s/n 72) prior to station 58 through the end of the cruise.

CTD Data Acquisition

The CTD data acquisition system consisted of an SBE-11plus (V2) deck unit s/n 111660 and a networked Dell Optiplex 755 PC workstation running Windows XP Professional. SBE Seasave v.7.21d software was used for data acquisition and to close bottles on the rosette.  Real-time digital data were backed up by the Data Manager, and post-cast raw data files were archived on a thumb drive as well as on Survey and PMEL networked PCs.  No real-time data were lost during this cruise.

CTD deployments were initiated by Survey after the Bridge advised that the ship was on station. The computer console operator maintained a CTD Cast log recording position and depth information at the surface, depth, and end of each cast; a record of every attempt to close a bottle, and any pertinent comments.

After the underwater package entered the water, the winch operator would lower it to a minimum of 10 meters and hold. The CTD pumps are configured with a 60-second startup delay, and were usually on by this time. The console operator checked the CTD data for reasonable values, waited an additional 60 seconds for sensors to stabilize, instructed the winch operator to bring the package to the surface, paused for 10 seconds, and descended to a target depth. The profiling rate was nominally 30 m/min to 50 m, 45 m/min to 200 m, and 60 m/min deeper than 200 m.  These rates could vary depending on sea cable tension and the sea state.

The console watch monitored the progress of the deployment and quality of the CTD data through interactive graphics and operational displays.  Additionally, the watch created a sample log for the cast that would later be used to record the correspondence between rosette bottles and analytical samples taken.  The altimeter channel, CTD pressure, wire-out, and bathymetric depth were all monitored to determine the distance of the package from the bottom, usually allowing a safe approach to within 10 meters.  The pinger was not turned on for the majority of casts owing to interference with the bathymetric depth.

Bottles were closed on the upcast via software, and were tripped 30 seconds after stopping at a bottle depth to allow the rosette wake to dissipate and the bottles to flush. The winch operator was instructed to proceed to the next bottle stop 10-20 seconds after closing bottles to ensure that stable CTD and reference temperature data were associated with the trip.

Near the surface, Survey directed the winch to stop the rosette just beneath the surface.  After the surface bottle was closed, the package was recovered.  Once on deck, the console operator terminated data acquisition, turned off the deck unit, and assisted with rosette sampling.

At the end of each cast, primary and secondary CTDO sensors were flushed with a solution of dilute Triton-X in de-ionized water using syringes fitted with Tygon tubing. The syringes were left attached to the temperature duct between casts, with the temperature and conductivity sensors immersed in the rinsing solution, to guard against airborne contaminants.

Acquisition Problems

The CTD was initially terminated on the aft .322 three-conductor winch cable.  A 400-dbar test cast was fully successful.  At station 1, cast 1 was bio-fouled during deployment, especially the secondary TCO sensors, and the cast was aborted at 170 dbar.  After the sensors were flushed repeatedly, cast 2 was acquired successfully.  Single modulo errors started at station 3 and increased to a few by station 10.  After a brake failure on the aft winch during station 10, the package was moved to the forward .322 three-conducting winch cable for two stations while it was repaired.

At station 9, cast 1 was bio-fouled at 20 dbar and aborted at 220 dbar.  During recovery the package was dropped less than a foot onto the deck/platform.  During cast 2, the primary conductivity was obviously bad and the secondary temperature was very noisy.  The cast was aborted at 600 dbar and these two sensors were replaced with s/n 2882 and s/n 4193.  It’s possible that the secondary conductivity cell was damaged at this time when the TC duct was stressed at an angle because during cast 3, the secondary conductivity was very noisy.  First, the secondary pump was replaced with s/n 5855 prior to station 10; and then the secondary conductivity was replaced with s/n 3858 prior to station 11 on the forward winch.

Modulo errors increased from 4 to 16 on the aft winch at stations 13-17.  Processed data had to be edited for spikes and bad data gaps in any one of the TCO data channels, including pressure.  While the aft winch was reterminated, stations 18-19 were done on the forward winch.  In addition to modulo errors, there were also “unsupported modem messages”.  The upcast at station 19 was done in two parts after cycling power on the deck unit and reinitializing software when bottle trips were no longer being confirmed.  Troubleshooting resulted in replacing the CTD prior to station 21.

A few modulo errors persisted occasionally on the aft termination but were accompanied by spikes and bad data gaps.  The aft winch developed controller problems after station 45 and was removed from service at station 51.  The forward winch was reterminated prior to station 51 and used for the remainder of the cruise in spite of level wind issues.

Prior to station 32, secondary oxygen sensor s/n 1890 was replaced because it had drifted too far from its calibration (over 15 umol/kg).  Prior to station 36, secondary oxygen sensor s/n 664 was replaced because its behavior was suspect.  Secondary oxygen data from s/n 2085 went bad for stations 102-103 and the sensor was replaced prior to station 104 with s/n 2040.

CTD Data Processing

The reduction of profile data began with a standard suite of processing modules using Sea-Bird Data Processing Version 7.21d software in the following order:

DATCNV converts raw data into engineering units and creates a .ROS bottle file.  Both down and up casts were processed for scan, elapsed time(s), pressure, t0, t1, c0, c1, oxvo1, oxvo2, ox1 and ox2.  Optical sensor data were converted to voltages and also carried through the processing stream.  MARKSCAN was used to skip over scans acquired on deck and while priming the system under water.

ALIGNCTD aligns temperature, conductivity, and oxygen measurements in time relative to pressure to ensure that derived parameters are made using measurements from the same parcel of water.  Primary and secondary conductivity were automatically advanced in the V2 deck unit by 0.073 seconds.  No further alignment was warranted.  It was not necessary to align temperature or oxygen.

BOTTLESUM averages burst data over an 8-second interval (+/- 4 seconds of the confirm bit) and derives both primary and secondary salinity, potential temperature, and potential density anomaly.  Primary and secondary oxygen were derived in DATCNV and averaged in BOTTLESUM, as recommended recently by Sea-Bird.

WILDEDIT makes two passes through the data in 100 scan bins.  The first pass flags points greater than 2 standard deviations; the second pass removes points greater than 20 standard deviations from the mean with the flagged points excluded.  Data were kept within 100 standard deviations of the mean (i.e. all data).

FILTER applies a low pass filter to pressure with a time constant of 0.15 seconds.  In order to produce zero phase (no time shift) the filter is first run forward through the file and then run backwards through the file.

CELLTM uses a recursive filter to remove conductivity cell thermal mass effects from measured conductivity.  In areas with steep temperature gradients the thermal mass correction is on the order of 0.005 PSS-78.  In other areas the correction is negligible.

Nominal values of 0.03 and 7.0 s were used for the thermal anomaly amplitude and the thermal anomaly time constant, respectively, as suggested by Sea-Bird.

LOOPEDIT removes scans associated with pressure slowdowns and reversals.  If the CTD velocity is less than 0.25 m s-1 or the pressure is not greater than the previous maximum scan, the scan is omitted.

DERIVE uses 1-dbar averaged pressure, temperature, and conductivity to compute primary and secondary salinity, as well as more accurate oxygen values.

BINAVG averages the data into 1-dbar bins.  Each bin is centered on an integer pressure value, e.g. the 1-dbar bin averages scans where pressure is between 0.5 dbar and 1.5 dbar.  There is no surface bin.  The number of points averaged in each bin is included in the data file.

STRIP removes oxygen that was derived in DATCNV.

TRANS converts the binary data file to ASCII format.

Package slowdowns and reversals owing to ship roll can move mixed water in tow to in front of the CTD sensors and create artificial density inversions and other artifacts.  In addition to Seasoft module LOOPEDIT, MATLAB program deloop.m computes values of density locally referenced between every 1 dbar of pressure to compute the square of the buoyancy frequency, N2, and linearly interpolates temperature, conductivity, and oxygen voltage over those records where N2 is less than or equal to -1 x 10-5 s-2.   Fourteen profiles failed the criteria in the top 4-10 dbars.  These data were retained by program deloop_post.m and were flagged as questionable in the final WOCE formatted files.

Program calctd.m reads the delooped data files and applies preliminary calibrations to temperature, conductivity, and oxygen; and computes calibrated salinity.

Pressure Calibration

Pressure calibrations for the CTD instruments used during this cruise were pre-cruise.  No additional adjustments were applied.  On deck pressure readings prior to each cast were examined and remained within 1 dbar of calibration.  Differences between first and last submerged pressures for each cast were also examined and the residual pressure offsets were less than 1 dbar.

Temperature Calibration

A viscous heating correction of -0.0006 C was applied at sea as recommended by Sea-Bird.  Post-cruise laboratory calibrations showed only slight changes in sensor behavior, including the SBE 35 reference temperature sensor.  Therefore it was not necessary to apply a drift correction to the data.

Conductivity Calibration

Seasoft module BOTTLESUM creates a sample file for each cast.  These files were appended using program sbecal.f.  Program addsal.f matched sample salinities to CTD salinities by station/sample number.

Each of the two primary conductivity sensors used were a single calibration grouping.  For conductivity sensor s/n 3157, program calcop0.m (a constant conductivity offset, a linear pressure-dependent correction to conductivity, and a constant conductivity slope produced the best fit to sample data from stations 1-8 for this primary sensor:

number of points used   143

total number of points  157

% of points used in fit 91.08

fit standard deviation  0.001534

fit bias                -0.00043480266

fit co pressure fudge   -2.2145704e-007

min fit slope           1.0000712

max fit slope           1.0000712

For conductivity sensor s/n 2882, program calcos1.m (a constant conductivity offset, and a 1st order polynomial conductivity slope as a function of station number) produced the best fit to sample data from stations 11-120 for this primary sensor:

number of points used   2218

total number of points  2552

% of points used in fit 86.91

fit standard deviation  0.001157

fit bias                 0.0011332164

min fit slope           0.99996306

max fit slope           0.99999932

Conductivity calibrations were applied to profile data using program calctd.m and to burst data using calclo.m.  CTD-bottle conductivity differences plotted against station number (Figure 1) and pressure (Figure 2) allow a visual assessment of the success of the fits.

Oxygen Calibration

A hybrid of the Owens-Millard (1985) and Murphy-Larson (revised 2010) oxygen sensor

modeling equations was used to calibrate the SBE-43 oxygen sensor data from this cruise.  The equation has the form

Ox=Soc*(V+Voff+Tau*exp(DI*P+D2*T).*dVdt).*Os.*exp(Tcor*T).*exp(Pcor*P./(273.15+T));

Where Ox is the CTD oxygen (in mol/kg), V is the measured oxygen voltage (in volts), dVdt is the temporal gradient of the oxygen voltage (in volts/s estimated by running linear fits made over 5 seconds), P is the CTD pressure (in dbar), T is the CTD temperature (in C), and Os is the oxygen saturation computed from the CTD data following Garcia & Gordon (1992).  Oxygen sensor hysteresis was improved by matching upcast bottle oxygen data to downcast CTD data by potential density anomalies referenced to the closest 1000-dbar interval using program match_sgn.m.  We used the values provided by SBE for each sensor for the constants D1 (1.9263e-4) and D2 (-4.6480e-2) to model the pressure and temperature dependence of the response time for the sensor. For each group of stations fit we determined values of Soc (sometimes station dependent), Voff, Tau, Tcor, and Pcor by minimizing the residuals between the bottle oxygen and CTD oxygen.  W represents fitting switches.  If the switches are set to 0,0 the fit is a regular L2 (least squares) norm for the entire group.  If the switches are set to 1,0 the fit is a regular L2 norm for the entire group but with a slope that is a linear function of station number.   If the switches are set to 2,0 the program first fits the entire group, then goes back and fits a slope and bias to individual stations, keeping the other parameters at the group values.  If the switches are set to 0,1 the fit is a regular L2 norma for the entire group but it is weighted by the nominal oxygen bottle spacing, thus fitting the deep portion of the water column better.

Program addoxy.f matched bottle sample oxygen values to CTD oxygen values by station/sample number. Program run_oxygen_cal_ml.m was used to determine calibration coefficients for eight station groupings determined by visual inspection:

Stns       Soc Range          Voff        Tau         Tcor       Pcor     Points Used    StdDev   W

1-15 0.4791-0.4853 -0.4718 7.230  0.0012 0.0398  324  92.3% 1.2012 1,0

16-61 0.4913-0.4963 -0.4706 8.395  0.0006 0.0391 1056  91.7% 1.6782 2,0

26        0.4769 -0.4588 4.826  0.0019 0.0422   21 100.0% 0.6601 0,0

54        0.4806 -0.4530 3.182  0.0019 0.0410   19 100.0% 0.4934 0,0

62-73        0.4971 -0.4699 8.730 -0.0001 0.0387  286  92.7% 1.0725 0,1

74-89        0.4838 -0.4504 8.698  0.0014 0.0389  377  94.4% 0.6334 0,0

90-97 0.4887-0.4850 -0.4583 8.330  0.0012 0.0392  191  99.5% 0.6356 1,0

98-120       0.4820 -0.4538 7.344  0.0014 0.0392  465  92.7% 0.6447 0,0

Oxygen calibration coefficients were applied to profile data using program calctd.m, and to burst data using calclo.m.

Primary sensor CTD - bottle oxygen differences plotted against station number (Figure 3) and pressure (Figure 4) allow a visual assessment of the success of the fits.

Despiking

Less than 8% of the profiles had to be despiked owing to electrical termination problems.  Single point spikes were interpolated by hand.  Salinity and oxygen were interpolated over larger ranges using program select_interp_ranges.m and apply_interp_sal_ox.m.  Interpolated records are indicated with WOCE quality flags of 6.