Chlorofluorocarbon and Sulfur Hexafluoride Measurements on the 2009 CLIVAR I5 Expedition

PIs:		John Bullister

Mark Warner

Analysts:	David Wisegarver

Eric Wisegarver

Erin Shields

Approximately 3500 samples were analyzed for two dissolved chlorofluorocarbons (CFC-11 and CFC-12) and for sulfur hexafluoride (SF6) on the CLIVAR I5 expedition, using methods described by Bullister and Wisegarver (2008).In general the analytical system performed well on the cruise.

Routine measurements of dissolved SF6 in seawater remain extremely challenging. Typical dissolved SF6 concentrations in modern surface water are ~1-2 fmol kg-1 seawater (1 fmol= femtomole = 10-15 moles), approximately 1000 times lower than dissolved CFC-11 and CFC-12 concentrations. The limits of detection for SF6 on CLIVAR I5 were approximately 0.02 fmol kg-1. SF6. Improvements in the analytical sensitivity to this compound at low concentrations are essential to make these measurements more routine on future CLIVAR cruises.

Water samples on CLIVAR I5 were collected in bottles designed with a modified end-cap to minimize the contact of the water sample with the end-cap O-rings after closing. Stainless steel springs covered with a nylon powder coat were substituted for the internal elastic tubing provided with standard Niskin bottles. When taken, water samples collected for dissolved CFC-11, CFC-12 and SF6 (‘CFC/SF6’) analysis were the first samples drawn from the bottles. Care was taken to coordinate the sampling of CFC/SF6 with other samples to minimize the time between the initial opening of each bottle and the completion of sample drawing. Samples most easily impacted by gas exchange (dissolved oxygen, 3He, DIC and pH) were collected within several minutes of the initial opening of each bottle. To minimize contact with air, the CFC/SF6 samples were drawn directly through the stopcocks of the bottles into 250 ml precision glass syringes equipped with three-way plastic stopcocks. The syringes were immersed in a holding tank of clean surface seawater held at ~10OC until ~20 minutes before being analyzed. At that time, the syringe was place in a bath of surface seawater heated to ~30oC.

For atmospheric sampling, a ~75 m length of 3/8" OD Dekaron tubing was run from the CFC van located on the fantail to the bow of the ship. A flow of air was drawn through this line into the main laboratory using an Air Cadet pump. The air was compressed in the pump, with the downstream pressure held at ~1.5 atm. using a backpressure regulator. A tee allowed a flow of ~100 ml min-1 of the compressed air to be directed to the gas sample valves of the CFC/SF6 analytical systems, while the bulk flow of the air (>7 l min-1) was vented through the back-pressure regulator. Air samples were analyzed only when the relative wind direction was within 60 degrees of the bow of the ship to reduce the possibility of shipboard contamination. Analysis of bow air was performed at 14 locations along the cruise track. At each location, at least five air measurements were made to increase the precision of the measurements. Air measurements are listed at the end of this report.

Concentrations of CFC-11, CFC-12 and SF6 in air samples, seawater, and gas standards were measured by shipboard electron capture gas chromatography (EC-GC) using techniques modified from those described by Bullister and Weiss (1988) and Bullister and Wisegarver (2008) as outlined below. For seawater analyses, water was transferred from a glass syringe to a glass-sparging chamber (volume ~200 ml). The dissolved gases in the seawater sample were extracted by passing a supply of CFC/SF6 free purge gas through the sparging chamber for a period of 6 minutes at ~150 ml min-1. Water vapor was removed from the purge gas during passage through an 18 cm long, 3/8" diameter glass tube packed with the desiccant magnesium perchlorate. The sample gases were concentrated on a cold-trap consisting of a 1/16" OD stainless steel tube with a 5 cm section packed tightly with Porapak Q (60-80 mesh) and a 22 cm section packed with Carboxen 1000. A Neslab Cryocool CC-100 was used to cool the trap to ~-70°C. After 6 minutes of purging, the trap was isolated, and it was heated electrically to ~200°C. The sample gases held in the trap were then injected onto a precolumn (~60 cm of 1/8" O.D. stainless steel tubing packed with 80-100 mesh Porasil B, held at 80°C) for the initial separation of CFC-12, CFC-11, SF6  and CCl4 from later eluting peaks.

After the SF6 and CFC-12 had passed from the pre-column and into the second precolumn (5 cm of 1/8" O.D. stainless steel tubing packed with MS5A, 80°C) and into the analytical column #1 (240 cm of 1/8" OD stainless steel tubing packed with MS5A and held at 80°C), the outflow from the first precolumn was diverted to the second analytical column (150 cm 1/8" OD stainless steel tubing packed with Carbograph 1AC, 80-100 mesh, held at 80°C). After CFC-11 had passed through the first pre-column, the flow was diverted to a third analytical column (1.7 m of Carbograph 1AC, 80°C). The gases remaining after CCl4 had passed through the first pre-column, were backflushed from the pre column and vented. Column #1 and the second pre-column were held in a Shimadzu GC8 gas chromatograph with an electron capture detector (ECD) held at 340oC. Column #2 and the first precolumn were in another Shimadzu GC8 gas chromatograph with ECD. Column #3 was held in a Shimadzu Mini2 gas chromatograph (90 C) with the ECD held at 250oC.

The analytical system was calibrated frequently using a standard gas of known CFC/SF6 composition. Gas sample loops of known volume were thoroughly flushed with standard gas and injected into the system. The temperature and pressure was recorded so that the amount of gas injected could be calculated. The procedures used to transfer the standard gas to the trap, precolumn, main chromatographic column, and ECD were similar to those used for analyzing water samples. Four sizes of gas sample loops were used. Multiple injections of these loop volumes could be made to allow the system to be calibrated over a relatively wide range of concentrations. Air samples and system blanks (injections of loops of CFC/SF6 free gas) were injected and analyzed in a similar manner. The typical analysis time for seawater, air, standard or blank samples was ~11 minutes. Concentrations of the CFC-11 and CFC-12 in air, seawater samples, and gas standards are reported relative to the SIO98 calibration scale (Cunnold et al., 2000). Concentrations of SF6 in air, seawater samples, and gas standards are reported relative to the SIO-05 calibration scale. Concentrations in air and standard gas are reported in units of mole fraction CFC in dry gas, and are typically in the parts per trillion (ppt) range. Dissolved CFC concentrations are given in units of picomoles per kilogram seawater (pmol kg-1) and SF6 concentrations in fmol kg-1. CFC/SF6 concentrations in air and seawater samples were determined by fitting their chromatographic peak areas to multi-point calibration curves, generated by injecting multiple sample loops of gas from a working standard (PMEL cylinder 45174) into the analytical instrument. The response of the detector to the range of moles of CFC/SF6 passing through the detector remained relatively constant during the cruise. Full-range calibration curves were run at intervals of 4-5 days during the cruise. Single injections of a fixed volume of standard gas at approximately one atm pressure were run much more frequently (at intervals of ~90 minutes) to monitor short-term changes in detector sensitivity.

The purging efficiency was estimated by re-purging a high-concentration water sample and measuring this residual signal.  At a flow rate of 150 cc min-1 for 6 minutes, the purging efficiency for all 3 gases was >99%.

On this expedition, based on the analysis of ~250 pairs of duplicate samples, we estimate precisions (1 standard deviation) of about 1% or 0.002 pmol kg-1 (whichever is greater) for both dissolved CFC-11 and CFC-12 measurements. The estimated precision for SF6 was 2% or 0.02 fmol kg-1, (whichever is greater). Overall accuracy of the measurements (a function of the absolute accuracy of the calibration gases, volumetric calibrations of the sample gas loops and purge chamber, errors in fits to the calibration curves and other factors) is estimated to be about 2% or 0.004 pmol kg-1 for CFC11 and CFC-12 and 4% or 0.04 fmol kg-1 for SF6).

A small number of water samples had anomalously high CFC/SF6 concentrations relative to adjacent samples. These samples occurred sporadically during the cruise and were not clearly associated with other features in the water column (e.g., anomalous dissolved oxygen, salinity, or temperature features). This suggests that these samples were probably contaminated with CFCs/SF6 during the sampling or analysis processes.

Measured concentrations for these anomalous samples are included in the data file, but are given a quality flag value of either 3 (questionable measurement) or 4 (bad measurement). Less than 2% of samples were flagged as bad or questionable during this voyage. A quality flag of 5 was assigned to water samples which were drawn from the rosette but lost during storage or due to errors in the multi-step analytical process.

References

Bullister, J.L., and R.F. Weiss, 1988: Determination of CC13F and CC12F2 in seawater and air. Deep-Sea Res., v. 25,  pp. 839-853.

Bullister, J.L., and D.P. Wisegarver (2008): The shipboard analysis of trace levels of sulfur hexafluoride, chlorofluorocarbon-11 and chlorofluorocarbon-12 in seawater. Deep-Sea Res. I, 55, 1063–1074.

Prinn, R.G., R.F. Weiss, P.J. Fraser, P.G. Simmonds, D.M. Cunnold, F.N. Alyea, S. O'Doherty, P. Salameh, B.R.  Miller, J. Huang, R.H.J. Wang, D.E. Hartley, C. Harth,  L.P. Steele, G. Sturrock, P.M. Midgley, and A. McCulloch, 2000: A history of chemically and radiatively important gases in air deduced from ALE/GAGE/AGAGE. J. Geophys. Res., 105, pp. 17,751-17,792.

Air Measurements on I05_2009

Concentrations are in pats-per-trillion (PPT)

Date		Time	CFC12	CFC11	SF6

YYMMDD	HHMM	PPT	PPT	PPT

090321	2335	532.7	239.3	6.39

090321	2342	534.0

090321	2349	532.0	240.2	6.49

090321	2356	531.8	241.9	6.45

090322	0003	531.3	244.2	6.52

090322	0010	531.7	243.5	6.55

090322	1412		238.4	6.52

090322	1419	532.3	237.2	6.49

090322	1426	536.4	237.3	6.45

090322	1433	536.1	237.8	6.37

090322	1440	534.1	238.0	6.37

090326	1003	536.4	240.6

090326	1010	536.2	240.5

090326	1017	532.3	240.0

090326	1024	532.8	240.5

090326	1031	533.2	241.2

090330	0135	535.6	240.8	6.70

090330	0145	530.5	240.4	6.77

090330	0155	531.5	239.9	6.69

090330	0205	529.1	240.1	6.69

090330	0215	532.2	239.9	6.56

090330	0225	532.0	240.4	6.70

090331	0706	528.6	239.5	6.34

090331	0713	525.2	239.5	6.37

090331	0720	535.7	240.2	6.61

090331	0727	528.1	239.9	6.52

090331	0734	530.9	240.6	6.50

090331	0741	534.1	239.8	6.51

090406	0050	532.2	240.6	6.46

090406	0057	530.8	240.2	6.43

090406	0104	531.6	240.9	6.51

090406	0111	531.1	240.2	6.67

090406	0118	534.2	240.5	6.63

090406	0125	535.4	240.3	6.53

090409	1825	531.0	240.3	6.63

090409	1835	531.7	239.3	6.53

090409	1845	535.2	239.4	6.59

090409	1855	530.1	239.7	6.64

090409	1905	532.2	239.7	6.68

090409	1915		240.2	6.64

090414	2145	532.6	239.9	6.45

090414	2155	533.5	239.0	6.72

090414	2205	534.7	240.0	6.67

090414	2215	531.8	238.7	6.58

090414	2225	531.8	239.5	6.68

090414	2235	532.4	239.2	6.67

090418	0524	529.2	237.0	6.56

090418	0534	529.8	235.3	6.65

090418	0544	531.5	233.9	6.72

090418	0554	533.5	232.0	6.72

090418	0604	529.3	231.0	6.75

090422	1143	533.8	244.3	6.43

090422	1153	535.2	245.4	6.40

090422	1203	534.4	244.8	6.44

090422	1213	535.0	245.2	6.51

090423	2057	533.5	242.3

090423	2107	534.5	242.7	6.38

090423	2117	532.5	243.3	6.34

090423	2127	533.6	244.2	6.38

090427	0335	531.6	240.1	6.81

090427	0346	533.2	238.7	6.74

090427	0357	527.7	239.0	6.66

090427	0408	530.9	239.0	6.51

090427	0419	531.4	241.4	6.39

090504	0450	539.0	238.7	6.46

090504	0501	537.7	236.0	6.65

090504	0512	537.3	236.6	6.61

090504	0523	539.3	238.1	6.68

090504	0534	543.4	239.1	6.72

090507	0559	532.4	240.1	6.69

090507	0611	530.4	240.4	6.76

090507	0623	529.9	239.9	6.71

090507	0635	532.1	241.2	6.68

090507	0647	532.2	239.5	6.70

Mean		532.8	239.8	6.57

STDEV		2.9	2.5	0.13

%STDEV	0.5	1.1	1.9