DISSOLVED INORGANIC CARBON (DIC) I09N 2025 (Fremantle to Phuket)

PIs: Richard A. Feely (Lead; NOAA-PMEL), Brendan Carter (UW-CICOES/NOAA-PMEL), and Rik Wanninkhof (NOAA-AOML)

Technicians: Abigail Tinari (UW-CICOES/NOAA-PMEL) and Chuck Featherstone (NOAA-AOML)

Final Data Submission Parameters (AT 11/24/2025)

Including the analytical replicates, 3102 samples were analyzed from 98 CTD casts for dissolved inorganic carbon (DIC) on the I09N 2025 GO-SHIP cruise. Excluding the Niskins that were dedicated for biological analysis (~15%), 84.4% of the remaining available Niskins were sampled. The accuracy of the DICE measurement is evaluated with the use of CRMs, supplied by Dr. A. Dickson of Scripps Institution of Oceanography (SIO). The DIC data reported have been corrected to the certified value (DIC = 2029.30 µmol/kg) for the batch (219) used. The summary table (Table 1) below lists information for the CRMs for each system. We collected replicate samples from approximately 10% of the Niskins and these analytical replicates were analyzed near the beginning, middle and end of each cell to serve as a measure of precision. The average absolute difference from the mean of these replicate measurements was 1.00 µmol/kg (0.92 µmol/kg (DICE 1) and 1.09 µmol/kg (DICE2); Table 2).

Table 1. CRM Certified value and system results

| CRM Info | PMEL1 |  |  | PMEL2 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Batch – Cert | Average | N | Std. Dev | Average | N | Std. Dev |
| 219 – 2029.30 | 2029.71 | 60 | 2.82 | 2026.97 | 70 | 2.5 |

Table 2. Calibration data during this cruise

| SYSTEM | Average Gas Loop Cal Factor | Pipette Volume | Average difference in duplicates |
| --- | --- | --- | --- |
| DICE 1 | 1.00741 | 27.603 | 0.92 |
| DICE 2 | 1.00448 | 26.403 | 1.09 |